## LOT 1 medical equipments
Annex 1 Technical Specifications
Item No. | Description | Unit | Quantity | Comment
Medical Equipment
1 | Bilirubinometer Filters: 461nm, 551nm, Range of measurement: 0 – 30mg/dL(Total bilirubin), Correcting Haemolysis: 0 – 250mg/dL Hbcv Measuring accuracy: ±5%, Sample Volume: 50 – 60μL, Display: 3-1/2, 7-segment red LED, Sample Container: Haematocrit capillary tube (heparinized), Light source: 6V 1.5A tungsten lamp, Detector: Silicon photocell: Power Requirements: 100 – 240VAC, 50/60Hz, 35W | pcs | 1
2 | Hot Plate Digital control • Overheat prevention circuit turns off the heater if the top plate temperature reaches 450°C • Hot top warning indicator • Microprocessor PID temperature control • Three user-selectable temperature control modes: - Optimal Mode - Fast Mode - Slow Mode • Two-user selective timer, showing actual and set value simultaneously • Non-slip heating bath (optional) • Transparent safety shield (optional) Temperature Range Ambient 50°C to 350°C Stirring Capacity (H20) : more than 4L Heating Control Feedback Control with PID Display LED digital (0.1C resolution) | pcs | 1
3 | Electrocardiograph (ECG) State of the art technology manufactured by world first class companies. 12 electrographic standard leads measuring with simultaneous acquisition. Visualization of at least one group of 3 leads simultaneously. At least n. 1 rest ECG patient cable with at least 5 terminals. TFT colour monitor display of at least 7 inches with visualization of analogical curves, alphanumeric values measured and the related physiological limits. Equipment compatible with patients with pacemakers. Protection against defibrillation. Automatic internal data storing up to at least 40 ECG records. Input impedance not less than 5MΩ. Patient leakage current no higher than 10uA. Common Mode Rejection (CMR) not less than +100dB. ECG signal measurement range not smaller than -2 mV to +2 mV. At least the following adjustable alarms: - heart failure; - ventricular fibrillation; - tachycardia; - bradycardia; - electrode disconnection. Equipment provided with filters at least for baseline instability, AC 50/60Hz interference, rumours, etc. Equipment provided with at least the following software applications: - arrhythmias detection and analysis; - ventricular fibrillation detection and analysis; - ventricular tachycardia detection and analysis; - ST segment analysis. Integrated printer. 12 lead ECG cable min. 100 sets of ECG connection electrodes (disposable type) N°6 suction ball-type reusable chest electrodes. N. 6 extremity clamp reusable electrodes. Cart/Trolley for the 12 Channel ECG Device. | pcs | 14
4 | Diagnostics Set Set Includes; Ophthalmoscope head, Otoscope head, set of three Long Aural Speculum, Set of four short Aural Speculum, Duplay Dilating Speculum, Throat Lamp Halogen, Laryngeal Mirror, Medium Battery Handle, Spare Halogen Bulb, Case |  | 1
5 | Oxygen Flowmeter Medical Oxygen Flow Meter for Bed Head Unit, with humidifier bottle & face mask material: Sterilizable Range: 0-10 litre | pcs | 135
6 | Bed Head Suction Unit Vacuum Regulator (Central Suction) Suitable for central vacuum Sterilizable With bottle and accessories Capacity: 2L, wall type Double hose connector Vacuum hose: adapter | pcs | 115
7 | Laryngoscope Laryngoscope set, Fibre Optic, LED Illumination, Mcintosh, 0, 1, 2, 3, 4 blades set, short handle, tale handle. - On/off switch to be robust and easy to use. - Autoclavable - Supplied in protective, enclosable container. | pcs | 15
8 | Nebulizer Ultrasonic type, for nebulization of liquids for inhalation therapy. Heavy duty and suitable for intensive use in a secondary level hospital. Latest generation, fully digital control panel and microcontroller technology. All the parts other than filters shall be autoclavable up to 132° Celsius. Supplied medicine caps allow the hospital staff to provide medicine to the patient efficiently. Water tank: min. 650 ml water tank with min. 1000 ml additional tank Adjustable nebulization for the adjustment of the given medicine amount in time. Accessories to be included: Trolley/mobile stand, hoses, filters, water bottle, bacterial filter, and air filter. All items in contact with the patient shall be autoclavable. Main Areas of usage: Paediatrics, Intensive Care, Internal medicine applications, lung diseases, Bronchial Asthma, Pneumonias, Acute and non-acute bronchial, Humidification of respiratory gases, Drug Nebulization. Main Features • Suitable for continuous use. Digital control panel for precise operation. • Built in timer, adjustable up to 90 minutes. • Automatic heating mechanism, heats vapour up to 38° C. • Constant vapor output • Alarm for malfunction and low water level • Volume of drugs to be applied: Min 5 ml –Max 7ml • Droplet/micron size should be less than 1.5mµ • Nebulization approx. 2ml/min. with no wastage of drug • Possibility of use combined with Oxygen therapy. • Oxygen adapter for supplying respiratory gases with the vapor. • Filtering system for reducing risk of contamination. • Controls: Intensity to control intensity of mist. Speed fan Min. 4; to control the volume of mist | pcs | 11
9 | Pulse Oximeter for adult and pediatric. Continuously displays patient oxygen saturation in real time using an external probe on the skin. Contains adjustable alarms to alert when either saturation or heart rate is low Reusable, sterilizable probes are robust and easily connected and disconnected Operates from mains voltage and from internal rechargeable battery Application: SPO2, pulse rate. Display- LED, Backlight illuminated. Parameters and waveform displayed -SPO2, pulse rate, system status, plethysmogram, SPO2 range 0-100 %. Audiovisual Alarms - high / low SPO2 and pulse rate, sensor off, sensor failure, low battery. Cable length should be minimum 1.5 meter. Interface for data communication. AC Power Cable. | pcs | 20
10 | Medical Electric plaster saw can effectively cut all kinds of bandages High-speed saw blade vibration system Cut all kinds of bandages Properties: Orthopedic Surgical Instrument three saw blades one wrench to load unload saw blades. | pcs | 1
11 | Infusion Pump Large volume infusion pump suitable for ICU, CCU, NICU, PICU, OR and Emergency Room. Suitable for adult and paediatric. User friendly by using intuitive interface and large arrow buttons. Easy-to-see screen displays in any lighting setting. • Entry drug library: approx. 1,000. • Higher flow rates: Up to approx.1200 ml/hr allowing for a wider range of therapies. Flow rate accuracy: ± 5 % • Alarms are audible, clearly visible and differentiated using different tones and colour coded alarms. • Clamp capable to holds up to four pumps. • Chargeable when not in use in the space station. • Battery pack (interchangeable) for about 8 hr. operation. • Hard, soft or both limits shall be possible to accommodate. • A pump should be small and lightweight. • Pumps shall require a dedicated IV pole. • Switches and controls should be protected against accidental setting changes. • The pump shall be protected against fluid spills due to breakage or disconnection of IV solution containers. Patient safety should not be compromised, and pump operation should not be adversely affected by fluid. | pcs | 1
12 | Syringe pump, single channel, AC and battery powered, with accessories Syringe pump, portable, for IV fluids infusion, single syringe Continuous delivery, piston driven Suitable for rail and IV pole mount Compatible syringes 10, 20, 30, 50 and 60 ml Automatic syringe capacity detection Open system, compatible with commercial syringe sets Programmable infusion by rate, time, flow and dose (patient body weight and drug concentration with automatic calculation) and with syringe capacity selection Flow-rate range, programmable: 0.1 – 400 ml/h for 10 ml syringe, 0.1 ml/h step 0.1 – 800 ml/h for 20 – 30 ml syringe 0.1 ml/h step 0.1 – 1,800 ml/h for 50 – 60 ml syringe 0.1 ml/h step Accuracy ± 2 % Volume limit range 0.1 – 999.9 ml, programmable, with 0.1 ml step Time limit range 1 – 1,999 min of operation Bolus program, maximum volume 25 ml or syringe capacity whichever higher Bolus range: 10 – 300 ml/h for 10 ml syringe 10 – 600 ml/h for 20 – 30 ml syringe 10 – 1,200 ml/h for 50 – 60 ml syringe KVO program, 1 ml/h Programmable infusion by patient body weight, drug concentration, with automatic calculations Occlusion detection, min. 3 selectable levels Display and control panel with visualisation of all working parameters (program, volume, flow, time, dose, patient data) and system messages Audio/visual alarms, with silencing feature, for: occlusion, plunger and | pcs | 37
syringe error, status and end of infusion/program (near empty, empty syringe), battery/power supply status and failure, system errors Automatic self-test function Battery and AC powered, with automatic switch between modes and automatic recharge when connected to mains Battery autonomy > 8 h (at 5 ml/h) Integrated battery charger Sturdy construction, splash- and shock-proof, plastic enclosure Power requirements 110 – 220 V / 50 – 60 Hz Power consumption 25 W including the following items: min. 200 disposable syringes each of volumes: 10, 20, 50, 60 and 100 ml with needles 1 x rail and IV pole mounting clamp 1 x spare battery pack Instructions for assembly, use and maintenance in English
13 | Suction Unit Aspirators Mobile suction pump, surgical with locking antistatic casters, Vacuum range: 0 to >500 mm Hg, - Flow rate: at least 25-40 litres per minute with accuracy of 10% Collection canister: 2 canisters - Capacity: ≥ 1500 mL each Vacuum control method: Regulator preferred. Overflow protection FILTER TYPE : Bacterial, hydrophobic Airline to pump to incorporate bacterial filter: Tubing to patient: to be minimum 3m long, non-collapsible type Easily visible control panel to include 'power on' indicator, vacuum control valve and vacuum gauge. Power Source: 230 VAC Hydrophobic bacterial filter" Recommended spares & consumables for the suction pump, such as - 40 disposable canisters. | pcs | 9
14 | Vein Viewer To detect subcutaneous blood and create a digital image of the patient’s vein pattern projected directly onto the skin’s surface in real time. Visualization Depth: Min. 10 mm to Max. 15mm, visualization as it happens to see peripheral veins, bifurcations and valves and assess in real time the refill/flushing of veins . Brightness, approx. : Min. 6 lumens . With HD imaging. Battery continuous run time ,Min. 2 hours . Power source : Lithium Ion Battery and AC . Imaging / focal distance : Approx. 33 cm . Base dimensions, approx. 45*60 cm with double caster break . Maximum arm extension, approx. : 60 cm to 65 cm . Height with the Head Unit down , approx. 150 cm | pcs | 1
15 | Tourniquet System, Automatic Automatic, Single for OR use, table or Pole Mount - Single Port - Dual Potential Cuff - Rechargeable Battery Approx. 6 hrs - Touchscreen Interface - Display: 7.0" Colour - Limb Occlusion Pressure - Line Occlusion Detection (Acoustic Reflectometry) - Cuff Alert - Automatic IVRA Cuff Lock Out - Advanced Leak Detection - Pre-Op Cuff Test | pcs | 3
16 | Orthopedic Drill blade mounting mechanism eliminates unwanted vibration and noise. Effortless attachment assembly, simply push and click. Simple blade loading Suitable for both large bone and trauma procedures Hand piece. Dual trigger Min. Drilling Speed: 0 - 1,350 rpm Min. Reaming Speed: 0 - 330 rpm Min. reaming torque: 15 Nm Cannulation diameter: 4.4 mm Dedicated Attachments: Min. 10 options One additional Rechargeable Battery Bone drill handpieces + Set with: - x2 Battery, with Battery Charger - Blade size: (0, 0.8, 1, 1.5, 1.8, 2, 2.3, 2.5, 2.7, 3, 3.2, 3.5, 4, 4.2, 4.5, 5, 5.4, 6, 7, 8)mm - Canulatet Drill Set of Saw handpieces: Reciprocating Saw, Sternum Saw, Sagittal Saw, Precision Saw - Sealed containers, Stainless steel sterilization cases, Ergonomic case handles protect against heat | pcs | 1
17 | Apron Lead for whole body and for Neck, Easy on / Easy off convenience; improved lower back support and an increased level of comfort and protection range from 0.35 to 0.5 mmpb Size : S,M,L,XL Set (1 set for each size) | pcs | 4
TOTAL
NOTES:
1. Unless otherwise explicitly requested, all standard accessories, maintenance tools/consumables/spare parts required for the proper start-up operation of the goods shall be included in the offer. Bidders shall specify optionally in the bid, the quantity and details of any such items Detailed technical specification, brochures, labelling and detailed information of the product as well as copies of international certificates (for each item, if necessary) for each product shall be attached to the bid.
2. Quality control for supplies: All goods must meet the requirements of manufacturing legislation and regulation of medical equipment in the country of origin; meet internationally recognized standards for safety, efficacy, and quality; conform to all the specifications and related documents contain herein: (i) be fit for the purposes expressly made known to the bidder by the buyer; (ii) be free from defects in workmanship and materials; and be certified by a competent authority in the manufacturer’s country.
3. Documents for Quality Assurance: The successful bidder might be requested to provide the buyer with a copy of quality certificate, certificate of origin, internationally recognized quality certificates, documents under which the quality assurance was processed in the manufacturer’s country, maintenance and repair manuals free of charge.
4. Necessity of voltage stabilizer to be mentioned in the offer
5. Wherever spare parts and/or consumables are not required during the first 36 months of equipment operation, the bidder shall state so in its bid. - Wherever installation is not required in the following detailed specifications, the supplier shall at least unpack the provided equipment, assemble the provided equipment, where required, change the wall socket, if so required, and test run the equipment in the presence of a representative of the buyer charged either with operating or maintaining the equipment. - Wherever user training is not required in the following detailed specifications, the supplier shall at least provide a brief introduction on the proper usage of the equipment.
6. Note: Concerning the hospital equipment to which training of the users, engineers and technicians at the end user’s hospital is needed it shall cover the following areas: - Introduction to the function of the medical equipment - Introduction to the manual of the equipment and the routine checks and maintenance. - Appropriate care of the equipment - Troubleshooting for simple failures and breakdowns - Calibration methods and necessity - Safety testing of medical equipment - Risk management and liability control

## Lot 2 medical Furniture
Annex 1 Technical Specifications
Item No. | Description | Unit | Qty | Comment
Medical and Non-medical Furniture
1 | Stool Height adjustment by gas spring Polyurethane legs with 5 swivel castors Foam seat with synthetic leather cover External diameter: approx. 60cm Height range: min. 56-71cm Seat diameter: approx. 42cm Castor diameter: approx. 50mm Safe working load capacity: min. 140kg | pcs | 23
2 | Strecher Type: Hydraulic Patient Transfer Trolley. Mattress: Anti-microbial, flame retardant, water proof and air permeable. Dimensions: (LxWxH): Approx. 2120 x 800x (630-930) mm. High adjustment range: Approx. +/- 55 to 90 mm Platform: Two sections, easily cleanable. C-Arm compatible with full length x-ray translucent platform. Height adjustment: Hydraulic Trendelenburg and reverse Trendelenburg: Approx. 0° – 12° (Degrees) Back rest adjustable angle: Approx. 0° – 70° (Degrees) Gas spring assisted back rest position. Safe load capacity: Approx. 220 kg. Side rails: 2pcs, collapsible. Pushing bars: Detachable. Castors: Four of approx. 200 mm in diameter castors with bowl-type central braking system, the unit shall be equipped with a 5th steer castor for easy maneuverability. Foot pedals for movements at both sides of patient stretcher. Four corners central lock pedal. I.V. drip pole sockets. Drainage bad holders and accessory hooks. Oxygen tank holder. | pcs | 8
3 | Bowl Stand, double Stainless steel main structure Two stainless steel bowl Height adjustable 360 degrees, the castors (tires) must be 360 degree of free movement Swivel and diagonal lockable castors Length: approx. 650mm Width: approx. 350mm Height range: approx. 795-1425mm Bowl diameter: approx. 260mm Bowl capacity: min. 10lt Castor Diameter: approx. 75mm | pcs | 14
4 | Operation Room Transfer Stretcher Manual connecting Transfer Stretcher, 2 Base + 1 Top, Cart A is inside the operation room while Cart B is outside the operating room, both carts are connected by adjusting them to the same height, Gas spring assisted backrest adjustment - Telescopic IV pole - Collapsible, full length side rails with safety lock mechanism - Easily cleanable, compact laminate two section mattress platform - 4 units of IV pole sockets and accessory sockets - Drainage bag holders and accessory hooks - Anti-microbial, water proof, flame retardant, air-permeable mattress - Detachable pushing bars - 360° (Degrees) Swivel and lockable castors Accessories: Oxygen tube holder for 2liter Safety belt Viscoelastic mattress Length: approx. 210cm External Length: approx. 415cm Width: approx. 70cm Backrest angle: approx. 0°-75° (Degrees) Height (without mattress): approx. 60cm Castor diameter: approx. 200mm Mattress surface: approx. 60x190cm Back section length: approx. 70cm Mattress thickness: Min. 7 – Max. 9 cm Safe working load capacity: approx 230kg | pcs | 2
5 | Medical Waste Trolley Heavy duty trolley, manufactured from 304/316 grade stainless steel or comparable with 4 swivel casters for effortless transportation, two lockable castors. Resistant to Hospital disinfecting solutions Leak proof construction and top covers. External Dimensions, (W × L × H): approx. 60 × 120 × 100 cm. | pcs | 6
6 | Laboratory test results pockets Wall mounted, consisted of Approx. 20-Pocket Vertical White Back and Clear Laboratory Test Result Material: Acrylic Each pocket holds approximately 20 test results Overall Dimension (W × H × D) Approx. 40" × 50" × 2" or 101 x 127 x 5 cm | pcs | 1
7 | Manual Blood Drawing Chair with armrests special for Blood Sampling and assure optimum comfort. Blood sampling chair with fixed height, Approx. 50 cm Comfortable upholstery with good density, width Approx. 50 x 50 cm Steel quality, epoxy paint Adjusting back rest via 1 Lockable Gas Spring Jack. Back Seat size Approx. 50 x 75 cm - Flexible from 95 to 175 Degrees. Equipped with 2 blood test splints size Approx. 40 x 15 x 10 cm, position adjustable by lock mechanism Load capacity: Approx. 150kg. | pcs | 1
8 | Multi Purpose Transport Trolley Heavy duty trolley for transporting sterile goods, made of micro-polished stainless steel. The doors with all-round silicone seals and a stop bar, help guarantee dust tight closure and prevent the contamination of the sterile products. Interior dimensions Approx. 1000 x 600 x 1100 mm - Two double-walled hinged doors with all-round seal and stop bar - Opening angle of doors: 270° (Degree) - Two vertical push handles on each side - Four corner guards - 2-point door lock - approx. Four galvanised heavy duty castors, two with brakes | pcs | 6
9 | X-Ray Apron Hangers Supporting three chrome plated Apron Hangers, consisting of a stainless steel backplate Approx. 110mm wide x 58mm high projecting Approx.165mm, complete with wall plugs and four fixing screws. | pcs | 2
10 | Closed Transport Trolley - To be put in the Sterile Store to OT - Size: approx. 1500x750x1260 mm - A trolley for sterile goods handling where higher than normal dust protection is required, e.g. short transports between hospital buildings. Suitable for handling baskets or containers with a total capacity of 9 STU (1 STU = 600 x 300 x 300 mm) on three solid, removable shelves (3 x 3 STU). - Trolley should be fitted with large stainless-steel wheels (Ø 160 mm) for easier manoeuvrability. - Two fixed wheels and two swivel wheels with brakes. - The fully welded stainless-steel construction (minimum 18 gauges, 304) makes it suitable for cabinet washers. The doors open 270° for easy access and cleaning. - Trolley should have lockable doors and should include handlebars. Supplier to perform installation, safety and operation checks before handover The case is to be cleanable with alcohol or chlorine wipes. Trolley for 5 full size sterile goods basket, with covers. Trolley to be used for transport between CSSD and storage area in O.T. Dept., and for storage in the preparation rooms in O.T. Dept. Rack trolley to be used for 5 full size wire baskets or 10 half size baskets the baskets are part of trolleys and therefore have matching dimensions Construction: Stainless steel tubes and guides, solid stainless steel bottom. Two s/s good sealed doors. Four castors, Two with brakes, washable. Disposable covers for trolley: 1000 pcs, covers are plastic bags (nylon) dimention approx 1600 x 800 x 1350 mm as part of first disposables | pcs | 6
TOTAL
NOTES:
1. Unless otherwise explicitly requested, all standard accessories, maintenance tools/consumables/spare parts required for the proper start-up operation of the goods shall be included in the offer. Bidders shall specify optionally in the bid, the quantity and details of any such items Detailed technical specification, brochures, labelling and detailed information of the product as well as copies of international certificates (for each item, if necessary) for each product shall be attached to the bid.
2. Quality control for supplies: All goods must meet the requirements of manufacturing legislation and regulation of medical equipment in the country of origin; meet internationally recognized standards for safety, efficacy, and quality; conform to all the specifications and related documents contain herein: (i) be fit for the purposes expressly made known to the bidder by the buyer; (ii) be free from defects in workmanship and materials; and be certified by a competent authority in the manufacturer’s country.
3. Documents for Quality Assurance: The successful bidder might be requested to provide the buyer with a copy of quality certificate, certificate of origin, internationally recognized quality certificates, documents under which the quality assurance was processed in the manufacturer’s country, maintenance and repair manuals free of charge.
4. Necessity of voltage stabilizer to be mentioned in the offer
5. Wherever spare parts and/or consumables are not required during the first 36 months of equipment operation, the bidder shall state so in its bid. - Wherever installation is not required in the following detailed specifications, the supplier shall at least unpack the provided equipment, assemble the provided equipment, where required, change the wall socket, if so required, and test run the equipment in the presence of a representative of the buyer charged either with operating or maintaining the equipment. - Wherever user training is not required in the following detailed specifications, the supplier shall at least provide a brief introduction on the proper usage of the equipment.
6. Note: Concerning the hospital equipment to which training of the users, engineers and technicians at the end user’s hospital is needed it shall cover the following areas: - Introduction to the function of the medical equipment - Introduction to the manual of the equipment and the routine checks and maintenance. - Appropriate care of the equipment - Troubleshooting for simple failures and breakdowns - Calibration methods and necessity - Safety testing of medical equipment - Risk management and liability control