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	UNOPS eSourcing v2021
	ITB Ref No: xxx



Section II: Schedule of Requirements
eSourcing reference: ITB/2021/34707
A. Summary of Requirements

Lot 1: Supply and Installation of Medical Equipment for Primary Health Centers in Anbar-Iraq
Lot 2.:Supply and Installation of Furniture for Primary Health Centers in Anbar-Iraq
Lot 3: COVID-19 Protection Supplies 

B. Lot 1: Technical specifications for Goods and Comparative Data Table- Medical Equipment 

	#
	UNOPS minimum technical requirements
	QTY
	Is bid compliant? Bidder to complete
	Details of goods offered. Bidder to complete

	1
	Supply and Test of Biochemistry Fuji film NX-500
	3
	☐ Yes   ☐ No
	Insert details of goods offered, including specifications and brand/model offered if applicable

	
	Application: 
Determine the concentration of certain metabolites, electrolytes, proteins, and/or drugs in samples of serum,  plasma, urine, cerebrospinal fluid, and/or other body fluids. 
	
	☐ Yes   ☐ No
	

	
	Configuration: Floor type or bench type
	
	☐ Yes   ☐ No
	

	
	Processing modes: Random access , Continuous
	
	☐ Yes   ☐ No
	

	
	Throughput, max tests/hr: 128 test 
	
	☐ Yes   ☐ No
	

	
	Sample Type: Serum/plasma , Urine , optional : CSF 
	
	☐ Yes   ☐ No
	

	
	Sample Size, µL: ≤ 100  
	
	☐ Yes   ☐ No
	

	
	Test Menu, min: To be specified by user  
	
	☐ Yes   ☐ No
	

	
	Programmed Tests: Yes
	
	☐ Yes   ☐ No
	

	
	User Definable Tests:: Yes 
	
	☐ Yes   ☐ No
	

	
	Method Used: End point, kinetic, ISE or other  
	
	☐ Yes   ☐ No
	

	
	Optical system: Photometric, Spectrophotometric or other 
	
	☐ Yes   ☐ No
	

	
	Light source: Halogen, Tungsten-halogen or other 
	
	☐ Yes   ☐ No
	

	
	Reagents Type : Liquid or solid 
	
	☐ Yes   ☐ No
	

	
	Substitution ( open or closed system ): Yes  
	
	☐ Yes   ☐ No
	

	
	Refrigerated onboard: yes if liquid  
	
	☐ Yes   ☐ No
	

	
	System Features
	
	☐ Yes   ☐ No
	

	
	Closed-tube sampling : No
	
	☐ Yes   ☐ No
	

	
	Direct sampling : Yes
	
	☐ Yes   ☐ No
	

	
	Liquid-level sensing : Yes
	
	☐ Yes   ☐ No
	

	
	Clot detection: Yes 
	
	☐ Yes   ☐ No
	

	
	Auto dilution : Yes
	
	☐ Yes   ☐ No
	

	
	Abnormal values flag : Yes 
	
	☐ Yes   ☐ No
	

	
	Auto verification: According to manufacture  
	
	☐ Yes   ☐ No
	

	
	Auto quality control : Yes  
	
	☐ Yes   ☐ No
	

	
	Auto calibration : Yes  
	
	☐ Yes   ☐ No
	

	
	Data Management:
	
	☐ Yes   ☐ No
	

	
	Display: Yes , touch screen 
	
	☐ Yes   ☐ No
	

	
	Results stored: Yes  
	
	☐ Yes   ☐ No
	

	
	Computer interface : Yes 
	
	☐ Yes   ☐ No
	

	
	Current LIS vendor interfaces : Yes  
	
	☐ Yes   ☐ No
	

	
	Printer: Yes
	
	☐ Yes   ☐ No
	

	
	Bar-code reader: Yes 
	
	☐ Yes   ☐ No
	

	
	Bar-code reader: Yes  
	
	☐ Yes   ☐ No
	

	
	LINE POWER, VAC, Hz: 220-240, 50/60 HZ  
	
	☐ Yes   ☐ No
	

	
	Backup or UPS: Yes , ≥ 30 min 
	
	☐ Yes   ☐ No
	

	
	water deionizer for supply: Yes 
	
	☐ Yes   ☐ No
	

	
	Onboard supply: Yes 
	
	☐ Yes   ☐ No
	

	
	Environmental requirements: The equipment suitable for working in the climate conditions of Iraq in terms of temperature  & humidity 
	
	☐ Yes   ☐ No
	

	2
	Supply and Test of CBC 3-Part Differential
	4
	☐ Yes   ☐ No
	Insert details of goods offered, including specifications and brand/model offered if applicable

	
	Application: Used to perform complete blood counts ,include WBC ,RBC ,Red blood cell and platelet count ,hemoglobin  concentration ,mean cell hemoglobin ( MCH), mean cell hemoglobin concentration (MCHC),mean cell volume (MCV)  
	
	☐ Yes   ☐ No
	

	
	Configuration: Bench type
	
	☐ Yes   ☐ No
	

	
	Intended area of use Laboratory  
	
	☐ Yes   ☐ No
	

	
	Automated/semi-automated: Fully automated 
	
	☐ Yes   ☐ No
	

	
	Method Used: Volumetric impedance  
	
	☐ Yes   ☐ No
	

	
	Test Menu
	
	☐ Yes   ☐ No
	

	
	Basic hematology: RBC, WBC, Hgb, Hct, MCV, MCH, MCHC ); Plt  
	
	☐ Yes   ☐ No
	

	
	WBC differentials: 3-part differential: L# and %, M# and %, G# and %  
	
	☐ Yes   ☐ No
	

	
	OTHER: RDW ,PDW, PCT & others tests can be performed by the device should be specified  
	
	☐ Yes   ☐ No
	

	
	Sample Type EDTA Whole blood required; Capillary or venous blood  
	
	☐ Yes   ☐ No
	

	
	SAMPLE VOLUME, µL  Start-up time, min ≤ 5  ≤ 120  
	
	☐ Yes   ☐ No
	

	
	THROUGHPUT, samples/hr ≥ 60  
	
	☐ Yes   ☐ No
	

	
	Analysis time, sec ≤ 60
	
	☐ Yes   ☐ No
	

	
	System Features:
	
	☐ Yes   ☐ No
	

	
	Auto dilution Yes  
	
	☐ Yes   ☐ No
	

	
	Auto-sampler Yes 
	
	☐ Yes   ☐ No
	

	
	Closed-tube sampling Yes
	
	☐ Yes   ☐ No
	

	
	Coincidence correct Yes
	
	☐ Yes   ☐ No
	

	
	Adjustable threshold Yes
	
	☐ Yes   ☐ No
	

	
	Histogram display Yes  
	
	☐ Yes   ☐ No
	

	
	Robotics capability NO  
	
	☐ Yes   ☐ No
	

	
	APERTURE (if the device don’t have aperture ) Number >1  
	
	☐ Yes   ☐ No
	

	
	SIZE (S), µm ≤100  
	
	☐ Yes   ☐ No
	

	
	Reagent Type Should be specified by manufacturer, Preparation No  
	
	☐ Yes   ☐ No
	

	
	Alert Indicators Yes ( should be specified in details )  
	
	☐ Yes   ☐ No
	

	
	Data Management
	
	☐ Yes   ☐ No
	

	
	Display, type LCD monitor , touch screen (preferred ) or other ( should be specified )  
	
	☐ Yes   ☐ No
	

	
	Data displayed Yes (results & other)  
	
	☐ Yes   ☐ No
	

	
	HIS/LIS interface Yes  
	
	☐ Yes   ☐ No
	

	
	Data entry Manual and Bar code   
	
	☐ Yes   ☐ No
	

	
	Data storage ≥ 10,000 patient results  
	
	☐ Yes   ☐ No
	

	
	Printer Yes  
	
	☐ Yes   ☐ No
	

	
	CALIBRATION Automatic  
	
	☐ Yes   ☐ No
	

	
	QUALITY CONTROL ≥ 3 files ,levey – Jennings charts , others should be specified 
	
	☐ Yes   ☐ No
	

	
	Line power, VAC 220-240, 50/60 HZ  
	
	☐ Yes   ☐ No
	

	
	Environmental requirements: The equipment suitable for working in the climate conditions of Iraq in terms of  temperature & humidity. 
	
	☐ Yes   ☐ No
	

	3
	Supply and Test of Centrifuge
	3
	☐ Yes   ☐ No
	Insert details of goods offered, including specifications and brand/model offered if applicable



	
	Spins 3ml-15ml test tubes at 5000 rpm 
	
	☐ Yes   ☐ No
	

	
	Minimum 16 tubes  
	
	☐ Yes   ☐ No
	

	
	Easy cleanout bowl  
	
	☐ Yes   ☐ No
	

	
	0 - 5000 RPM  
	
	☐ Yes   ☐ No
	

	
	Digital programmable control and readout 
	
	☐ Yes   ☐ No
	

	
	Variable speed and time settings 
	
	☐ Yes   ☐ No
	

	
	Lid safety lock  
	
	☐ Yes   ☐ No
	

	
	high g-forces for fast separation  
	
	☐ Yes   ☐ No
	

	
	Rubber feet to prevent slipping  
	
	☐ Yes   ☐ No
	

	
	Impact-resistant lid  
	
	☐ Yes   ☐ No
	

	
	Zero-rpm interlocking lid  
	
	☐ Yes   ☐ No
	

	
	Wide mouth for easy clean-up  
	
	☐ Yes   ☐ No
	

	
	Metal rotor and tube shields  
	
	☐ Yes   ☐ No
	

	
	Brushless motor  
	
	☐ Yes   ☐ No
	

	
	Angled and (swinging )tube holder as alternative  
	
	☐ Yes   ☐ No
	

	
	Fast breaking  
	
	☐ Yes   ☐ No
	

	
	Programmable time and speed  
	
	☐ Yes   ☐ No
	

	
	Rubber feet to prevent slipping
	
	☐ Yes   ☐ No
	

	
	POWER REQUIREMENTS :220v ,50 Hz  
	
	☐ Yes   ☐ No
	

	
	Application : Used to perform complete blood counts ,include WBC ,RBC ,Red blood cell and platelet count ,hemoglobin  concentration ,mean cell hemoglobin (MCH), mean cell hemoglobin concentration (MCHC),mean cell volume (MCV) 
	
	☐ Yes   ☐ No
	

	
	CONFIGURATION : Bench top  
	
	☐ Yes   ☐ No
	

	
	Intended area of use : Laboratory
	
	☐ Yes   ☐ No
	

	
	Automated/semi-automated: Fully automated  
	
	☐ Yes   ☐ No
	

	
	METHOD USED : Volumetric impedance  
	
	☐ Yes   ☐ No
	

	
	Test Menu
	
	☐ Yes   ☐ No
	

	
	Basic hematology : RBC, WBC, Hgb, Hct, MCV, MCH, MCHC ); Plt  
	
	☐ Yes   ☐ No
	

	
	WBC differentials : 3-part differential: L# and %, M# and %, G# and %  
	
	☐ Yes   ☐ No
	

	
	OTHER : RDW ,PDW, PCT & others tests can be performed by the device should be specified  
	
	☐ Yes   ☐ No
	

	
	SAMPLE TYPE : EDTA Whole blood required; Capillary or venous blood  
	
	☐ Yes   ☐ No
	

	
	SAMPLE VOLUME, µL: ≤ 120  
	
	☐ Yes   ☐ No
	

	
	THROUGHPUT, samples/hr : ≥ 60  
	
	☐ Yes   ☐ No
	

	
	Analysis time, sec ≤ 60  
	
	☐ Yes   ☐ No
	

	
	Start-up time, min ≤ 5  
	
	☐ Yes   ☐ No
	

	
	System Features
	
	☐ Yes   ☐ No
	

	
	Auto dilution Yes  
	
	☐ Yes   ☐ No
	

	
	Auto sampler Yes  
	
	☐ Yes   ☐ No
	

	
	Closed-tube sampling YES  
	
	☐ Yes   ☐ No
	

	
	Coincidence correct YES  
	
	☐ Yes   ☐ No
	

	
	Adjustable threshold YES  
	
	☐ Yes   ☐ No
	

	
	Histogram display yes  
	
	☐ Yes   ☐ No
	

	
	Robotics capability NO 
	
	☐ Yes   ☐ No
	

	
	APERTURE (if the device don’t have aperture ) 
	
	☐ Yes   ☐ No
	

	
	Number >1 
	
	☐ Yes   ☐ No
	

	
	SIZE (S), µm ≤100  
	
	☐ Yes   ☐ No
	

	
	REAGENT TYPE Should be specified by manufacturer  
	
	☐ Yes   ☐ No
	

	
	Preparation No  
	
	☐ Yes   ☐ No
	

	
	ALERT INDICATORS Yes ( should be specified in details )  
	
	☐ Yes   ☐ No
	

	
	Data Management
	
	☐ Yes   ☐ No
	

	
	Display, type LCD monitor , touch screen (preferred ) or other (should be specified)
	
	☐ Yes   ☐ No
	

	
	Data displayed Yes (results & other)  
	
	☐ Yes   ☐ No
	

	
	HIS/LIS interface Yes  
	
	☐ Yes   ☐ No
	

	
	Data entry Manual and Bar code  
	
	☐ Yes   ☐ No
	

	
	Data storage ≥ 10,000 patient results  
	
	☐ Yes   ☐ No
	

	
	Printer Yes  
	
	☐ Yes   ☐ No
	

	
	CALIBRATION Automatic  
	
	☐ Yes   ☐ No
	

	
	QUALITY CONTROL ≥ 3 files ,levey – Jennings charts , others should be specified  
	
	☐ Yes   ☐ No
	

	
	Line power, VAC 220-240, 50/60 HZ  
	
	☐ Yes   ☐ No
	

	
	Environmental requirements: The equipment suitable for working in the climate conditions of Iraq in terms of  temperature & humidity.
	
	☐ Yes   ☐ No
	

	4
	Supply and Test of Microscope
	3
	☐ Yes   ☐ No
	Insert details of goods offered, including specifications and brand/model offered if applicable

	
	Purpose : to provide magnification and resolution of specimens under examination
	
	☐ Yes   ☐ No
	

	
	Heavy duty and compatible design  
	
	☐ Yes   ☐ No
	

	
	OBSERVATION TUBES Type : binocular  
	
	☐ Yes   ☐ No
	

	
	Eyepieces:, 10x,15x, wide field, high-eye point  
	
	☐ Yes   ☐ No
	

	
	Interpupillary distance adjustment; mm: ≥( 55 -75),Adjustable.  
	
	☐ Yes   ☐ No
	

	
	NOSEPIECE CONFIGURATION: Quadruple or more 
	
	☐ Yes   ☐ No
	

	
	Objectives Magnification: (4x,10x,40x,100x) ,other should be specified  
	
	☐ Yes   ☐ No
	

	
	Objectives Type : Achromatic, plan achromatic, plan apochromatic ,or others  
	
	☐ Yes   ☐ No
	

	
	TOTAL MAGNIFICATION: ≥(40x-1000x)  
	
	☐ Yes   ☐ No
	

	
	Contrast Method
	
	☐ Yes   ☐ No
	

	
	Bright field: yes  
	
	☐ Yes   ☐ No
	

	
	Dark field: yes  
	
	☐ Yes   ☐ No
	

	
	Phase: optional  
	
	☐ Yes   ☐ No
	

	
	ILLUMINATION  
	
	☐ Yes   ☐ No
	

	
	Condenser type: Abbe, achromatic &/ or other type  
	
	☐ Yes   ☐ No
	

	
	Light source: LED (preferred),halogen, or other  
	
	☐ Yes   ☐ No
	

	
	STAND  
	
	☐ Yes   ☐ No
	

	
	Focusing mechanism: Moving stage, rack and pinion  
	
	☐ Yes   ☐ No
	

	
	Coarse, fine adjustments: Coaxial, low mounted, ≥ 1µ increments  
	
	☐ Yes   ☐ No
	

	
	STAGE Type: Mechanical or better  
	
	☐ Yes   ☐ No
	

	
	Tension adjustment: Coaxial  
	
	☐ Yes   ☐ No
	

	
	Stage Motion: X-Y  
	
	☐ Yes   ☐ No
	

	
	Numerical aperture: 0.9 to 1.4  
	
	☐ Yes   ☐ No
	

	
	Dust cover: yes
	
	☐ Yes   ☐ No
	

	
	Environmental requirements: the equipment suitable for working in the climate conditions of Iraq in terms of temperature &  humidity
	
	☐ Yes   ☐ No
	

	
	POWER SUPPLY: 220-240V,50-60Hz
	
	☐ Yes   ☐ No
	

	5
	Supply and Installation of Water Bath
	2
	☐ Yes   ☐ No
	Insert details of goods offered, including specifications and brand/model offered if applicable

	
	Applications: sample thawing, bacteriological examinations, warming reagents, coliform  determinations and microbiological assays  
	
	☐ Yes   ☐ No
	

	
	Heavy duty and compatible design  
	
	☐ Yes   ☐ No
	

	
	Material: seamless stainless steel  
	
	☐ Yes   ☐ No
	

	
	External: stainless steel, powder coated  
	
	☐ Yes   ☐ No
	

	
	Temperature range: 1ºC above ambient to 100ºC  
	
	☐ Yes   ☐ No
	

	
	On off switch and display by thermometer graded for temperature  
	
	☐ Yes   ☐ No
	

	
	Accuracy: 0.2ºC   
	
	☐ Yes   ☐ No
	

	
	Temperature preset buttons 
	
	☐ Yes   ☐ No
	

	
	Circulation Pump to mix water
	
	☐ Yes   ☐ No
	

	
	Microprocessor controlled  
	
	☐ Yes   ☐ No
	

	
	Heater: 1000- 1400 watts 
	
	☐ Yes   ☐ No
	

	
	Safety Device for overheating protection  
	
	☐ Yes   ☐ No
	

	
	Bath Volume: 8 liters  
	
	☐ Yes   ☐ No
	

	
	With stainless steel gabled cover  
	
	☐ Yes   ☐ No
	

	
	The exterior should remain cool, glass wool insulator  
	
	☐ Yes   ☐ No
	

	
	Alarm indicator for overheat limits, low water cutoff  
	
	☐ Yes   ☐ No
	

	
	Power Requirements 220-240 V, 50/60 Hz)
	
	☐ Yes   ☐ No
	

	6
	Supply & Test of Dental Unit & Chair
	2
	☐ Yes   ☐ No
	Insert details of goods offered, including specifications and brand/model offered if applicable

	
	Application: - Used for diagnostic and treatment of teeth and other application. 
	
	☐ Yes   ☐ No
	

	
	Pneumatic or Electric dental unit.  
	
	☐ Yes   ☐ No
	

	
	Gearbox motor or hydraulic for dental chair.  
	
	☐ Yes   ☐ No
	

	
	Minimum height of dental chair (35 ~ 38cm).  
	
	☐ Yes   ☐ No
	

	
	Maximum height of dental chair (75 ~ 80 cm)  
	
	☐ Yes   ☐ No
	

	
	Adjustable headrest.  
	
	☐ Yes   ☐ No
	

	
	Light projector with arm.  
	
	☐ Yes   ☐ No
	

	
	Dentist tray:  
	
	☐ Yes   ☐ No
	

	
	Control panel (with X-ray viewer). 
	
	☐ Yes   ☐ No
	

	
	3-way syringe.  
	
	☐ Yes   ☐ No
	

	
	3-module Midwest (4-hole).  
	
	☐ Yes   ☐ No
	

	
	Ultrasonic scaler.  
	
	☐ Yes   ☐ No
	

	
	2 turbines (Midwest 4 hole, standard head, stainless steel body, ceramic bearing, push button chuck & autoclave able) one with  fiber optic and small head and other turbine standard head without fiber optic.
	
	☐ Yes   ☐ No
	

	
	Air motor (forward & reverse control)  
	
	☐ Yes   ☐ No
	

	
	Contra angle (1:1 direct drive, push 
	
	☐ Yes   ☐ No
	

	
	button chuck, ball bearing, autoclave able, , non-optic).  
	
	☐ Yes   ☐ No
	

	
	Straight hand piece (1:1 direct drive, non-optic). 
	
	☐ Yes   ☐ No
	

	
	Hand piece-hanging type.  
	
	☐ Yes   ☐ No
	







B. Lot 2: Technical specifications for Goods and Comparative Data Table- Furniture 

	#
	UNOPS minimum technical requirements
	Unit
	QTY
	Is bid compliant? Bidder to complete
	Details of goods offered. Bidder to complete

	1
	Supply and in install metal a good brand  Locker with half Glass door with all accessories needed the supplier should submit for approval before supplying process 
	Ea
	63
	☐ Yes   ☐ No
	Insert details of goods offered, including specifications and brand/model offered if applicable

	2
	Supply and in install metal a good brand  Waiting Area  Seats (each set should contain 8 seats at least ) with all accessories needed the supplier should submit for approval before supplying process 
	Set
	33
	☐ Yes   ☐ No
	Insert details of goods offered, including specifications and brand/model offered if applicable

	3
	Supply and in install wood from  a good brand  Office Table dim (1.8mx0.7m x1.2m height ) with drawers and  all accessories needed the Contractor should submit for approval before supplying process 
	Ea
	58
	☐ Yes   ☐ No
	Insert details of goods offered, including specifications and brand/model offered if applicable

	4
	Supply and in install metal a good brand  Single chair with arms  with Leather and all accessories needed the Contractor should submit for approval before supplying process 
	Ea
	58
	☐ Yes   ☐ No
	Insert details of goods offered, including specifications and brand/model offered if applicable

	5
	Supply and install aluminum Counter with dim (90 cm H, 60 cm D) using for Lab room. 
	L.M
	35
	☐ Yes   ☐ No
	Insert details of goods offered, including specifications and brand/model offered if applicable

	6
	Supply and install aluminum reception Counter Table with Aluminum Partition (4 m * 1.2 m * 0.6 m) 
	Ea
	1
	☐ Yes   ☐ No
	Insert details of goods offered, including specifications and brand/model offered if applicable











B. Lot 3: Technical specifications for Goods and Comparative Data Table- covid 19- Protection Supplies:

	#
	UNOPS minimum technical requirements
	Unit
	QTY
	Is bid compliant? Bidder to complete
	Details of goods offered. Bidder to complete

	1
	Supply three layers disposable Masks each package containing 50 PCS. 
	Ea
	800
	☐ Yes   ☐ No
	Insert details of goods offered, including specifications and brand/model offered if applicable

	2
	Supply disposable gloves type white Latex Rubber Gloves (XL)  each package contain 50 pcs
	Ea
	400
	☐ Yes   ☐ No
	Insert details of goods offered, including specifications and brand/model offered if applicable

	3
	Supply instant  hand Sanitizers spray ,alcohol 70% with Moisturizers & Vitamin E,size 300 ml.e for each bottle 
	Ea
	375
	☐ Yes   ☐ No
	Insert details of goods offered, including specifications and brand/model offered if applicable

	4
	Disposable Full Body Protective Apron in Blue with waist sting and thumb loops each package contain 60 pcs 
	Package
	70
	☐ Yes   ☐ No
	Insert details of goods offered, including specifications and brand/model offered if applicable

	5
	Go Hooked 350 Microns Safety Face Shield, Anti-fog Full Face Shield, (Pack of 2)
	Ea
	375
	☐ Yes   ☐ No
	Insert details of goods offered, including specifications and brand/model offered if applicable


















C. Delivery requirements and Comparative Data Table

	UNOPS Requirements
	Is bid compliant? Bidder to complete
	Details 
Bidder to complete

	Delivery time
	The Bidder shall deliver all the items within 8 calendar weeks from the date of placing the purchase order by UNOPS.
	☐ Yes   ☐ No
	Insert details 

	Delivery place and Incoterms rules
	DAP - Location of Delivery: As per the distribution list for the seven health centers in Iraq as per Incoterms 2020
	☐ Yes   ☐ No
	Insert details 

	Consignee details
	TBC
	☐ Yes   ☐ No
	Insert details 

	UNOPS Right to vary requirements
	At the time the Contract is awarded, UNOPS reserves the right to vary the quantity of the goods and associated services specified above, provided this does not exceed +/- 20% without any change in the unit prices or other terms and conditions of the ITB.
	☐ Yes   ☐ No
	Insert details 




D.  Inspections and tests 

The following inspections and tests shall be performed: 

· UNOPS or its representative may inspect and/or test any or all items of the goods to confirm their conformity to the contract, prior to dispatch from the manufacturer’s premises. Such inspection and clearance will not prejudice the right of the consignee to inspect and test the goods on receipt at destination.

· If the goods fail to meet the laid down specifications, the supplier shall take immediate steps to remedy the deficiency or replace the defective goods to the satisfaction of the purchaser.
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