## Fayda Lab BOQ
ANNEX 1 TECHNICAL SPECIFICATION
Supply of Laboratory equipment for Fayda Water Treatment Plant
No | Description | Unit | Qty
1 | Drying Oven Capacity of 55 lt Operation Range: temperature 5Cº-250 Cº digital control panel 230V 50Hz | pcs | 1
2 | Digital Balance Precision:0.0001g Glass Cabinet LCD screen | pcs | 1
3 | Turbidity Meter (Bench Unit) Nephelometric Measurement Range 0.1-1000 NTU Protection Class:IP67 240 V | pcs | 1
4 | Conductivity meter (Bench Unit) with 5070 cell for general use stirrer,stand,buffer and calibration beaker set | pcs | 1
5 | PH meter ( Bench unit) Digital with probes | pcs | 1
6 | Spectrophotometer wavelength range:190-1100 mn wavelength resolution:0.1 nm wavelength accuracy:-+1 mn,50-60 Hz | pcs | 1
7 | Incubator Volume:55 lt Voltage: 230V Frequency:50 Hz Temprature Operation Range: Ambient Temperature 5-80 Cº Temperature sensor :Pt 100 Control System: PID Temperature change :+-0.1Cº Temperature Sensitivity: +-0.1 Cº | pcs | 1
8 | Autoclave Vertical type autoclaves which allow for top loading are suitable for high volume sterilization Chamber Volume: 50 L Number of baskets: 2 or more and 1 exhaust tank stainless inlet tank Sterilization Temperature 110-123 ºC Safety: Over temperature protector Voltage: 230 V Frequency: 50 Hz | pcs | 1
9 | water bath suitable for testing (E.coli) test temperature 44 Cº including test tube holder range (10-250)rpm | pcs | 2
10 | Bunsen Burner with well finished metal surfaces H: approx. 15 cm Inlet: approx. 8mm | pcs | 2
11 | Automatic Pipette Filler OmniPET or equivalent Suitable for 0.5 ml to 100 ml pipette Ergonomically shaped handle Sensitive valves for precise work with low volume pipettes Protected by filter and safety valve Autoclavable nosepiece and pipette holder | pcs | 4
12 | Refrigerator Type: Two Door Capacity: Approx. 400 L , compressor cooled. Frost : Yes Installation type : Free standing Temperature control adjustable. | pcs | 1
13 | Bottletop Dispenser for dispensing aggressive reagents autoclavable fast volume adjustment with analog slide 1-10 mL | pcs | 3
14 | Jar test capacity :1000 ml 10-300 rpm illumated | pcs | 1
15 | laboratory fume hood Biobase FH 1500 or equivalent External dimensions: 1500x840x2150 mm Internal dimensions: 1380x730x745 mm Air Velocity: 0.3 - 0.8 m/sec System Exhaust Volume: 990 m3/hr LED Lamp 16W*1 UV Lamp: 20W, emission of 253.7 nanometers for mot efficient decontamination front window motorized, 5mm toughened glass, anti-UV Exterior Body cold rolled steel with anti bacterial powder coating and work table with chemical resistant phenolic resin 2 waterproof sockets power supply: AC 220V +- 10%, 50/60 Hz; 110V+-10%, 60 Hz consumption 500 W supply includes water tap, gas tap, water sink, base stand, UV lamp, LED lamp, sockets, exhaust duct, active carbon filter, foot switch, HEPA filter | pcs | 3
16 | digital termometer with hour,humidity and degree features in ºC | pcs | 2
17 | Water Distiller Distillation Capacity 4 lt/h Water Consumption: 1 lt/min Conductivity at 20ºC: 1.5 µS/cm Measures Height x Width x Depth mm: 435 x 370 x 220 Heaters: 3000 W Electricity: 230 V. - 50/60 Hz. 13 A | pcs | 1
18 | Comparator Supplied field-ready in a carrying case with reagents and accessories for 30 tests Wide range of parameters for a variety of applications Color discs have a continuous color gradient making it possible to achieve a direct color match - no need to extrapolate the result between two color standards Dimensions: 98 x 54 x 162mm Display: Large LCD display with backlight Interfaces: Micro-USB External Storage: Micro SD card Measuring Principle: Optical DO Auto - OFF: Yes Portability: Portable Environmental Conditions: -5 - 50 Degree C Compliance: CE Protection Class: IP 67 | pcs | 2
19 | Hot Plate laboratory water heater and magnetic stirrer for 5 L stirring volume and adjustable temperature & speed control can heat and stirr at the same time Speed Range: 100 – 1,500 rpm Temperature Range: 50 – 500° C Power Consumption: 1020 Watts Heat Output: 1,000 W Dimensions Set Up Plate: 200 x 200 mm | pcs | 1
20 | Digital timer Display Type: LCD Range: 99min 59 sec Countdown and count up timerLoud alarm at the completion of preset time Portable and handy | pcs | 1
21 | Cuvette for Spectrophotometer UV Quartz cuvette light path 10 mm Volume: 3.5 mL outer dimensions: 12.5 x 12.5 x 45 mm | pcs | 2
23 | Pipette Stands Circular Type, holding up more than 30 pipettes | pcs | 2
24 | Desiccator 300 mm Diameter Inner: 25 cm Diameter Outer: 28 cm Height: 33 cm Product Type: Top Loading Vacuum glass Silicone O-ring ensures an airtight seal | pcs | 2
25 | Filter paper set material: cellulose sheets feature: Grade 1 pack of 100 parameter: 0.25 psi wet burst 150 sec/100 mL speed (Herzberg) W × L: 460 mm × 570 mm thickness: 180 μm ash: ≤0.06% pore size: 11 μm (Particle retention) basis weight: 87 g/m2 | pcs | 1
26 | Glassware | SET | 1
| Dispensmate 10ml with 1L bottle | pcs | 1
| Pressmatic dispensmate 10ml with 0.5L bottle | pcs | 3
| Digital burette with 1L brawn bottle | pcs | 2
| Graguated cylinder 1000 ml pack 2 | pack | 1
| Graguated cylinder 500 ml pack 10 | pack | 1
| Graguated cylinder 250 ml pack 10 | pack | 1
| Graguated cylinder 100 ml pack 10 | pack | 1
| Graguated cylinder 50 ml pack 10 | pack | 1
| Beaker Pyrex 1000ml pack 10 | pack | 1
| Beaker Pyrex 500ml pack 10 | pack | 1
| Beaker Pyrex 250ml pack 10 | pack | 1
| Beaker Pyrex 100ml pack 10 | pack | 1
| Beaker Pyrex 50 ml pack 10 | pack | 1
| Volumetric Flasks, Pyrex glass conical narrow neck 1000ml pack 2 | pack | 2
| Volumetric Flasks, Pyrex glass conical narrow neck 500ml pack 2 | pack | 2
| Volumetric Flasks, Pyrex glass conical narrow neck 250ml pack 5 | pack | 1
| Volumetric Flasks, Pyrex glass conical narrow neck 100ml pack 5 | pack | 2
| Volumetric Flasks, Pyrex glass conical narrow neck 50ml pack 5 | pack | 2
| Conical Flasks narrow neck graduated 250 pack 10 | pack | 1
| Funnel 12cm Ø | pcs | 3
| Reagent bottle Pyrex glass 100ml | pcs | 20
| Laboratory Porcelain Mortar with Pestle(510 ml) | pcs | 2
| Watch glass Pyrex 100 ml | pcs | 5
| test tube 180 * 18 mm with caps | pcs | 50
| Nessler tube 50 ml | pcs | 10
| pipette rack ,20-30 opening | pcs | 5
| Durham Tubes 50*6mm round bottom _rimless | pcs | 2
| Petri Dish Pyrex 100*20mm | pcs | 10
| Graduated Pipett Pyrex 10ml pack 10 | pcs | 2
| Pipette bulb one mark glass 100ml pack 2 | pack | 4
| Pipette bulb one mark glass 50ml pack 5 | pack | 1
| Pipette bulb one mark glass 25ml pack 5 | pack | 1
| sterilizer box for petri dishes | pcs | 2
27 | Chemicals | SET | 1
| HCL 32% 2.5L | pcs | 1
| Ammonium hydroxide 32% 2.5L | pcs | 1
| Murixide (pack/25g) | pcs | 1
| Isopropyl alcohol (2-Propanol) | pcs | 1
| sodiumthiosulfate | kg | 1
| EDTA | kg | 1
| Ammonium chloride | kg | 1
| Silver nitrate (pack/25g) | pcs | 1
| Eriochrome black T (pack/25g) | pcs | 1
| Methyl red (pack/25g) | pcs | 1
| Bromocresole green (pack/25g) | pcs | 1
| sodium hydroxide | kg | 1
| Maconcky broth | kg | 3
| sodium chloride | kg | 2
| Potassium chromate(pack/100g) | pcs | 1
| Calcium Carbonate | kg | 1
| HCL 1:1 | pcs | 1
| Stock Aluminum or Aluminum Potassium Sulfate ALK(SO4)2 .12H2O (500 gm) | pcs | 1
| Sulfuric acid 0.02N | pcs | 1
| Ascorbic Acid 250 gm | pcs | 1
| Sudium Acetate (NaC2H3O2.3H2O) 500gm | pcs | 1
| Acetic Acid 2.5 LT | pcs | 1
| Eriochrome Cyanin R 100 gm | pcs | 1
| Cadmium 100 gm | pcs | 1
28 | TOOLS | set | 1
| Lab Beaker Tongs | pcs | 4
| Paper filter F140 grade , 120mm Ø , pack. 100unit | pcs | 5
| Bunsen Burner | pcs | 2
| Tripod Stands | pcs | 1
| pipette filler | pcs | 2