Contracting Owner Delegate : La Chaîne de l’Espoir

Call for Tender 3

Tender Rules & Technical Specifications

Call for Tender 3

Tender Rules & Technical Specifications

Medical equipment tender

for the Sinjar French Medical Centre

Lot 1: Imaging equipment

Lot 2: Monitoring, Ventilation equipment

Medical equipment tender

for the Sinjar French Medical Centre

Lot 1: Imaging equipment

Lot 2: Monitoring, Ventilation equipment

The object of the consultation:

Closing date and time:

Wednesday 18th Oct,2023. 5:30 PM (Iraq time GMT+3)

Wednesday 18th Oct,2023. 5:30 PM (Iraq time GMT+3)

Table of Contents

TABLE OF CONTENTS	2

I.	INTRODUCTION TO LA CHAINE DE L’ESPOIR (CDE)	4

II.	OBJECT OF THE TENDER – MAIN CHARACTERISTICS	5

II.	1 Project context	5

II.	2 The object of the tender	5

III.	BIDDING CONDITIONS AND INFORMATION	6

III.1.	Scope of the consultation	6

III.2.	Variants	6

III.3.	Period of validity of the offers	6

III.4.	Indicative Calendar	7

III.5.	Tender rules & Technical specifications content	7

III.6.	Acknowledgment	8

III.7.	Request for clarification	8

III.8.	Your Offer	8

III.9.	Remuneration	13

III.10.	Bid evaluation	13

III.11.	Submission of proposals	13

III.12.	Bid acceptance	15

III.13.	General conditions	16

IV.	GENERAL ARTICLES	17

MANUFACTURER’S AUTHORIZATION FORM	18

V.	LIST OF ANNEXES:	22

VI.	TECHNICAL SPECIFICATIONS	23

VI.1.	Lot 1: Supply, installation, qualification and start-up of Imaging Equipment.	23

VI.2.	Lot 2: Supply, installation, qualification and start-up of Monitoring and Ventilation equipment.	29

All Tenderers are kindly advised to note the followings:

The ability of suppliers to provide on-site after-sales service in Sinjar at SFMC , either independently or through one of their official representatives, will be an important criterion in the contract awarding process.

- In submitting a tender, the Tenderer accepts in full and without restriction the special and general conditions governing this request for proposals and the proposed contract, added in Annex 12, as the sole basis of this tendering procedure, whatever his conditions of sale may be, which he hereby waives. Tenderers are expected to examine carefully and comply with all instructions, forms, contract provisions, and specifications contained in this tender dossier.

- It is strongly recommended that the tender rules document and its annexes be read thoroughly. Failure to observe the procedures laid out therein may result in disqualification from the evaluation process.

# INTRODUCTION TO LA CHAINE DE L’ESPOIR (CDE)

La Chaîne de l’Espoir is an international NGO founded in 1994. Its mission is to strengthen healthcare systems to give everyone, especially children, the same chances of survival and development.

Helping the most marginalized children is more than an ethical duty, it is a vital requirement for achieving the United Nations 2030 goals for children's health and well-being.

We work in 27 countries with a global approach, focusing on children, mothers and disadvantaged communities:

  - Prevention and screening from an early age, particularly through school health programs.

  - Care and surgery to meet the most urgent needs of children and their mothers.

  - Training and transfer of skills to local teams thanks to our international network of experts in all areas of surgery.

  - Construction	and	equipment	of	hospital	facilities	adapted	to	local	needs.

Within the framework of its missions, La Chaîne de l'Espoir and its Infrastructure and Biomedical department strive to meet a dual objective:

  - Giving the most vulnerable populations access to quality medical and surgical care through the construction or rehabilitation of medical structures.

  - Contribute to the strengthening of local medical and technical capacities through the implementation of training courses delivered by La Chaîne de l’Espoir experts.

La Chaîne de l'Espoir works in countries where medical infrastructures are lacking or limited in their human and financial means. The establishments are often poorly equipped, with high medical costs that do not allow for appropriate treatment of patients and their pathologies.

For further information on CDE, its mandate, and operations please see: https://www.chainedelespoir.org/en

# OBJECT OF THE TENDER – MAIN CHARACTERISTICS

## 1 Project context

Sinjar City, in the north-eastern mountains of the Iraqi Nineveh Governorate, lost most of its population and infrastructures under Daesh, including its hospital and primary care centers. To restore this territory and encourage the return of refugees, the Nadia’s Initiative NGO (from the 2018 Nobel peace prize laureate Nadia Murad) promoted rebuilding projects such as the Sinjar Hospital. With the support of the Iraqi Ministry of Health (MOH), as well as the financial and technical support of the French Government through its Ministry of Foreign Affairs - Crisis and Support Center (CDCS), CDE has been mandated as Contracted Owner Delegate for the future public hospital Sinjar French Medical Centre (SFMC).

The future hospital will:

  - Allow inhabitants and returnees of the Sinjar District to find a reply to their health needs,

  - Start rebuilding the future with a hospital for all.

The conception phase and execution drawings have been done with international standards, with French and Lebanese design offices.

The construction phase is under progress with an Iraqi Construction Contractor and followed by a Lebanese Construction Manager. The works on site have begun on December 6th 2021 and the planned Taking-Over Date is on Dec 16th 2023.

The access to medical contractors for installing the equipment will begin at the beginning of 2024.

## 2 The object of the tender

The present call of tender relates to soliciting bids of medical equipment of the SFMC, which are planned to be installed on-site from Feb 2024, according to the biomedical tender schedule as precise in Annex 6. CDE invites qualified Suppliers to make their firm best offers.

The Suppliers must respond without any restrictions to the technical specifications mentioned in the tender dossier.

The tender is divided into 2 independent lots. The Suppliers must respond in 2 independent offers for the 2 lots, which may be awarded independently, one from each other.

The tender is divided into 2 independent lots (categories):

- Lot 1: Imaging equipment

- Lot 2: Monitoring, Ventilation equipment

##### The Supplier may respond to one OR the two lots.

Therefore, any offer for each independent lot not answering all the above requirements within this lot will be rejected.

Project location: Sinjar City, Iraq

The quality of the products and the quality of their implementation will be decisive in achieving the objectives set.

Tenderers must strictly adhere to all requirements of the Tender Rules. No changes, substitutions, or other alternations to the rules and provisions stipulated in this note may be made or assumed unless it is instructed or approved in writing by CDE. Submission of Proposals shall be deemed as an acknowledgment by the Tenderers that all obligations stipulated by this Tender Rules will be met and unless specified otherwise, the Tenderers have read, understood, and agreed to all the instructions.

CDE has a policy of zero tolerance on prescribed practices, including fraud, corruption, collusion, unethical practices, and obstructions. Any proposal submitted will be regarded as an offer by the Tenderers and does not constitute or imply the acceptance of any Proposal by CDE. CDE is under no obligation to award a contract to any Tenderers as a result of this tender. The Ethics Clauses are given in Annex 8 and have to be signed. The non-respect of these clauses leads to breach of contract.

# BIDDING CONDITIONS AND INFORMATION

## Scope of the consultation

This restricted tender allows the contracting authority (CDE) to hold discussions with the applicant businesses (Tenderers/suppliers) in order to identify and select the best offer, based on principles of equity, and transparency, in order to guarantee the medical equipment of the SFMC according to standards defined in the tender technical specifications.

## Variants

The proposal of variants is authorized. All variants will be presented to CDE for final decision, according to the contract (Annex: will be provided by CDE later).

All materials, equipment or any item shall follow the recommended brands and origins.

## Period of validity of the offers

The period of validity of the offers is set at 180 days from the tendering closing date. CDE will make its best effort to award the contract within this period. The pricing model quoted in the Supplier’s proposal will remain valid for the duration of the Contract.

## Indicative Calendar

This calendar is indicative and could be changed by CDE at any moment. The Tenderers can’t ask for any compensation in case of changes.

| Activity | DATE | TIME (Iraq time) |
| --- | --- | --- |
| Publications | Monday, 24Sep 2023 | 9:00 am |
| Deadline for request for any clarifications from the Contracting Authority | Thursday 5th Oct,2023 | 5:30 pm |
| Last date on which clarifications are issued by the Contracting Authority | Thursday, 12th Oct,2023 | 5:30 pm |
| Deadline for submission of tenders | Wednesday, 18th Oct,2023 | 5:30 pm |
| Opening Meeting | Thursday 26th, Oct 2023 | 2:00 pm |
| Clarifications and negotiation. | Monday, 13th Nov, 2023 | 4:00 pm |
| Appointment of the selected suppliers & start of contract’s negotiation. | Thursday,07 Dec,2023 |  |
| Signature of the medical equipment contracts | Beginning of January, 2024 |  |

## Tender rules & Technical specifications content

-Tender rules & Technical specifications

-Annex 1Pricing Schedules for lot 1 Imaging equipment (To be completed and submitted)

- Annex 2 Pricing Schedules for lot 2 Monitoring and Ventilation equipment (To be completed and submitted)

Annex 3 to 5 : not applicable.

Annex 6: Biomedical tender schedule (To be signed and submitted) Document Annex 7: SFMC presentation and layouts

Annex 8: CDE Ethics clauses (To be signed and submitted)

Annex9: Impartiality document (To be signed and submitted)

Annex  10: Non-disclosure agreement (To be signed and submitted)

Annex 11: Checklist Tender documents submission (To be signed and submitted)

Annex 12 : Supply contract template

## Acknowledgment

We would appreciate you informing us of the receipt of this tender, as precise above, by return e-mail to: biomedtender@chainedelespoir.org, ONLY.

To:

    - Confirm the receipt of this tender by signing the complete list, attached in the email, with your signature and the date,

    - Whether or not you will be submitting a bid.

## Request for clarification

Tenderers are required to submit any requests in English only for clarification in respect of the Tender Rules by e-mail only to: biomedtender@chainedelespoir.org, ONLY.

Tenderers are requested to keep all questions concise and to submit them before 5th Oct,2023 before 5:30 pm

In compliance with CDE’s procedures, all Tenderers will be treated equally (Collective answers are given to all questions asked by one Tenderer).

Site visits: Tenderers are encouraged to visit the site during working hours to have a clearer idea of the site location, and its constraint, as logistical-wise.

The site visit is on his own responsibility and the Tenderer will be considered as fully aware of all constraints and dispositions on site.

The Tenderer acknowledges that it has been informed about the construction works in progress, that any site visit or any delivery on site will be possible before CDE’s written validation, and that the first installations should be planned after the Taking-Over of the Construction Site, in complete coordination with the Construction Contractor and the Construction Manager, who will be still on site too.

## Your Offer

- YOUR OFFER SHALL BE PREPARED IN ENGLISH (All supporting documents should also be in English)

- Offers that don’t respect the tender rules may be rejected.

- Please submit your offer using the Annexes (1,2) financial offer provided.

- Your offer shall comprise the following three sets of documents:

  - Administrative documents

  - Technical offer

  - Financial offer

###### Content of the administrative documents

Local Suppliers

A company or scientific bureau should provide the following valid documentation:

- Administrative Orders & Articles of association for the current year 2023 in Federal Iraq.

- Company Registration documents: If you are Scientific bureau:

  - Pharmacist syndicates license.

  - Import license.

  - Valid Tax ID/ clearance.

  - Recent Confirmation letter from Pharmacist syndicates registration.

  - Bank account details. If you are Medical company:

  - Ministry of trade: registration license.

  - Chamber of commerce.

  - Import license.

  - Valid Tax ID/ clearance.

  - Bank account details.

  - Recent Ministry of Health / Directorate General of Health approval of registration.

- Company classification: The Tenderer shall provide a valid document stating its classification/grade

- Experience of Firm: For each project, bidders must provide a good performance certificate. Details of the list of the projects with the certificates of client satisfaction should also be submitted, as well as a recommendation letter

International suppliers

International suppliers are required to submit equivalent documents in accordance with the applicable laws of their respective countries. Additionally, they must provide evidence demonstrating their legal capacity to offer on-site after-sales service in Sinjar at SFMC North Iraq, either through independent means or by utilizing one of their authorized representatives

##### Important Note: Failure to send the above-requested information/documents may result in disqualification of your proposal from further evaluation.

###### Content of the Technical offer

Important Note: No pricing information should be included in the technical proposal. The technical proposal should contain all information required.

The technical component of your proposal should be concisely presented and structured as indicated below, the following elements shall be provided in the Technical Offer:

- Understanding of the requirements: The Tenderer shall provide an executive summary that defines the overall approach to manage and operate all of the required works. The offer therefore should clearly show a consistent approach and timeframe through:

  - Detailed delivery, installation and training schedule. Please clearly state the completion period for each lot.

  - Precisions regarding the origin, the technical specifications of the main materials and the references of the corresponding suppliers.

  - A summary note indicating the main measures planned to ensure safety on- site, respect of the labor law, and prevention against child work.

  - A summary note indicating the measures taken regarding environment protection (waste management, resource protection, …).

  - A summary note indicating his capacity to work in Sinjar during all the installation and warranties period, and his feedback about the site visit done by his own.

- Maintenance and operation of equipment precise the plan for the preventive and curative maintenance for the equipment especially with specific contract or team available for the follow-up on site.

- Copy of the Leaflets / Technical Brochures / Product Data Sheets of the Model offered highlighting features in support of the information provided in TECHNICAL SPECIFICATION for desired LOT.

- Copy of Quality Certificates ISO, FDA, CE for each offered equipment

- Manufacturer’s Authorization letter

- Declaration of warranty

- Room layout with medical equipment installed and installation drawings

###### Content of the Financial Offer

The following elements shall be provided in the Financial Offer:

- Price offer: The Tenderer shall provide a Financial Offer containing prices quoted in USD against each line item given in the Pricing Schedule Form. The Financial offer is to be submitted as per the Pricing Schedule Forms (Annex 1,2). Bids that have a different price structure may not be accepted.

Any missing items can be added to the Financial Offer, as well as any Variant, with the mention “Variant” in the description line.

##### Company proforma invoice for quoted items with break up prices.

The Contract Price is a Lump-sum price.

- Your financial offer must be all-inclusive including but not limited to the cost of materials, labor, overheads, overtime, risks factors e.g., inflation, changes in customs duties, etc.

- The financial must to include the years guarantee for all equipment installed on site, it will start after the signature of all the required documents: Delivery Release Form, Installation Release Form, Commissioning Release Form, Training Release Form and Final Hand-Over sign by parties

- The cost of preparing a bid and of negotiating a contract, including any possible related travel, is not reimbursable nor can it be included as a direct cost of the assignment.

- The Financial Offer shall be submitted both in Excel and in a signed PDF.

- Insurances and Guarantees: The Tenderer has to contract all insurances and guarantees mentioned in the Contract for the Provision of the Works, and to send us a copy of their third-party civil liability policy,

- Ethics clauses / Impartiality document / non-disclosure agreement: The Tenderer has to read and sign these documents provided in Annexes 8, 9 and 10.

## Remuneration

All details regarding the payment of the medical equipment will be specified in the Contract

(Annex : will be provided by CDE later)

CDE shall then pay the Contract Price to the Supplier in the form of milestones in arrears against completed Works.

## Bid evaluation

- Administrative Evaluations:  Submitting the documents mentioned in (Article 111.8.4.a), the registration documents, company classifications

- Technical evaluation: quality and pertinence of the submitted documents, conformity to the requirements mentioned Article (111.8.4.b) and (Article IV)and (a) encompass the following:

  - Company profile: such as (company number of branches, technical team capacity (number of engineers), Similar project in Iraq.

  - Technical criteria: include equipment specification, after-sale services (installation, delivery time, training and warranty) and periodic maintenance

- Financial evaluation: global price, coherence, Proforma Invoice, Warranty, and precision of the price offer, conformity to the requirements mentioned in ( 111.8.4.b)  )mandatory

Each proposal from a Supplier will be considered separately and independently. Tenderers shall submit a complete proposal for each solicitation in which they wish to participate. References to previous or ongoing proposals will not be considered. Award of a previous contract with CDE will not be considered in itself as a preference or guarantee for the award of future solicitations on the same subject.

## Submission of proposals

- By email ONLY, preferably using a OneDrive link to have all documents together:

Technical and financial offer to: biomedtender@chainedelespoir.org, ONLY.

Please state the following in your e-mail subject field: “Biomed tender – Lot [Number of lots] - [Name of your firm] - [Serial number of the e-mails]” (example: 1/3, 2/3, 3/4).

E.g., an offer from Company “ABC” who is sending a total of 2 emails will be titled as follows: “Biomed tender – Lot [Number of lots] – ABC – 1/2”.

Please be aware of the fact that the e-mail policy employed by CDE limits the size of each e- mail to a maximum of 20 MB so it may be necessary to send more than one e-mail for the technical proposal.

The Financial Proposal can be submitted as a separate e-mail. The file name must bear the following reference “Biomed tender – FINANCES - Lot [Number of lots] - [Name of your firm] - [Serial number of the e-mails]”.

All attachments should be in PDF format, and Excel sheet too for the Financial Offer. All documents in your proposal should be clearly labeled.

Upon submission of your bid via email, you should receive an auto-reply confirming receipt. If you do not receive the auto-reply, please, inform us before the closing date of the tender.

##### Deadline of Submission: 18th Oct, 2023 5:30 (Iraq time)

Any bid received after this date or sent to another CDE address may be rejected. CDE may, at its discretion, extend the deadline for the submission of bids, by notifying all concerned.

CDE will not be responsible for locating or securing any information that is not identified in the bid. Accordingly, to ensure that sufficient information is available, the Tenderer shall furnish, as part of the bid, any descriptive material such as extracts, descriptions, and other necessary information it deems would enhance the comprehension of its offer. CDE may ask for clarifications regarding the submitted information and may also ask for the submission of additional documents if required.

Inclusion of copies of your offer with any correspondence sent directly to the attention of the responsible buyer or any other CDE staff other than the submission e-mail address is forbidden. Please send your bid directly to the address provided in Article B “By email” section K of this tender rules.

## Bid acceptance

Please note that CDE is not bound to select any of the firms submitting bids and does not bind itself in any way to select the firm offering the lowest price. Furthermore, the contract will be awarded to the bid considered most responsive to the needs.

CDE has the right to cancel the tender in case of Force Majeure, changes to the overall project, major incident during the process, or lack of qualitative answers, without the Tenderers being allowed to ask for any financial compensation.

CDE reserves the right to accept the whole or part of your bid.

CDE may, at its discretion, increase or decrease the proposed quantity when awarding the contract and would not expect a significant variation in the rate submitted. Any such increase or decrease in the contract duration would be negotiated with the successful Tenderer as part of the finalization of the contract.

CDE may, at its discretion, extend the deadline for the submission of bids, by notifying all prospective Suppliers in writing. The extension of the deadline may accompany a modification of the solicitation documents prepared by CDE at its initiative or in response to a clarification requested by a prospective Supplier.

Please note that CDE strictly follows zero tolerance policy and as such, advises its Suppliers not to offer any gifts, favor, hospitality and commission etc. to CDE staff. Any Supplier found to be offering gifts, favor, hospitality and commission etc. to CDE staff will be placed on the sanctioned list and CDE will not do business with them anymore.

## General conditions

Scope: CDE invites tenders for the works described and summarized by procedures, conditions and contract terms, as described in the above tender terms. CDE reserves the right to vary the quantity of work/materials specified in the Tender Documents without any changes in unit price or other terms and conditions and to accept or reject any, all, or part of submitted offers.

Performance Standards: The Work will be completed according to all applicable national and international standards, regulations, and statutes currently applicable to Iraq.

Environmental Standards: Suppliers should, at minimum, comply with all statutory and other legal requirements relating to the environmental impacts of their business. Areas which should be considered are, waste management, conservation, energy consumption, and overall environmental and social stainability.

The Contract provided in Annex (which will be sent later): Suppliers shall accept the contract provided without any major modification.

##### Tender Basis:

- The working language of this tender and Application is English, in case of any contradiction between the two translations and Arabic in the bill of the quantities, the English language version will prevail.

- All applications shall be made in accordance with the Tender Documents.

- All Applicants will receive identical documents: No applicant should add, omit, or change any item, term or condition on original papers.

- If Applicants have any additional requests and conditions, these shall be stipulated in a separate letter accompanying their application.

- Each applicant may make one application only.

- Each application shall be valid for the period of 180 days from its date of submission.

- Applications shall be made in writing, with prices calculated in USD, and be inclusive of All taxes, compulsory payments, levies, and duties.

- Any requests for clarifications regarding the project that are not addressed in written documents must be presented to CDE through email. The answer to any question raised in writing by any applicant will be issued to that applicant by email. It is a condition of this tender that no clarification shall be deemed to supersede, contradict, add to or detract from the conditions hereof, unless made in writing and signed by CDE or its designated representative.

##### CDE has a right to refuse the bids, if include the following:

- Are not registered companies.

- Are bankrupt or in the process of going bankrupt.

- Have been convicted for an offense concerning professional conduct.

- Have been guilty of grave professional misconduct (proven by any means which the Contracting authorities can justify).

- Have not fulfilled obligations related to payment of social security and taxes.

- Are guilty of serious misinterpretation in supplying information.

- Are in situations of conflict or potential conflict of interest with CDE.

- Were declared at the serious fault of implementation owing to a breach of their contractual Obligations.

### GENERAL ARTICLES

#### ARTICLE 1: General requirements

- The bids must include all the equipment necessary for the correct operation of the device.

- All material will be according to the international high quality

- The implementation will therefore be done with the optimum level to ensure the correct implementation of the installations, and the proper functioning from the start-up.

- The Tenderer may offer additional equipment and accessories not expressly requested in these specifications (or tender) if he deems them useful and suitable to satisfy the buyer.

#### ARTICLE 2: Compliance with texts

- All supplies must comply with the texts and current standard at the time of installation and commissioning.

- The supplier is required to meet the regulatory requirements for placing on the market in accordance with laws of medical devices and decrees of the Public Health Code.

- The attestation certificates must be attached to the Tenderer's proposal, specifying the class of the device as well as the appendices that served as a reference

#### ARTICLE 3: Manufacturer’s authorization Form

The authorization certificate must be attached to the Tenderer's proposal includes the following format:

### MANUFACTURER’S AUTHORIZATION FORM

(to be submitted by authorized manufacturer/importers in a letterhead)

No. Dated:

To : La Chaîne de l’Espoir

SFMC/ MADAM- Medical equipment tender n°1 committee Dear Sir / Madam,

Equipment/Item Name:

- We ……………………………….………………… (Name of the OEM) are the original manufacturers	of	the	above	equipment	having	a	registered	office	at

…………………………………………………………………………………. (Full address with telephone number/fax number & email ID and website), having factories at

and		,	do	hereby	authorize

M/s.	............... (Name and address of Tenderer) as submit bids, and subsequently negotiate, sign the contract and to supply, install & commission the above items with you against the above bid no……………………………….

- We also hereby undertake to provide full guarantee/warranty /CMC/AMC as agreed by the Tenderer in the event the Tenderer is changed as the dealers or the Tenderer fails to provide satisfactory after sales and service during such period of Comprehensive warranty/CMC/AMC and to supply all the spares/reagents / consumables for 5 years.

- We also hereby declare that we have the capacity to manufacture and supply, install and commission the quantity of the equipment’s bided within the stipulated time.

(Name)

for and on behalf of M/s.	 Date: (Name of manufacturers)

Place:

Seal

Note: This letter of authority should be on the letterhead of the manufacturing concern and should be signed by a person competent and has the power of attorney to bind the manufacturer. In case the distributor is quoting through the importer, then the manufacturer has to give authorization to importer and the importer has to give the authorization to the distributor in the above format.

#### ARTICLE 4: Training

The offer must include full support: travel, accommodation, training costs, meals (lunch, accommodation) for the following personnel:

- End user (user training): The medical and paramedical team (day and night team).

- The Biomedical service (technical training): The biomedical engineer and/or the biomedical maintenance technicians (at least 3 persons)

- Technical training must be completed after the installation.

- This will be sanctioned by a personal certificate, issued to each participant, stipulating that the level of skill required to ensure maintenance has been acquired by the technician.

During this training, all the necessary information must be provided:

- Complementary technical documentation.

- Set of computer access codes to the maintenance software, if applicable.

- List of the main spare parts to be kept in stock for the hospital

- The overall training offer must indicate the training schedule (periods, times) as well as the place and the planned training methods.

Reminder:

The Tenderer also undertakes to train again, at his own expense, the user, medical and biomedical teams, in the event of changes to the devices supplied (hardware or software changes).

The training may be postponed of several months after the installation and commissioning phases, if the biomedical and technical teams of the SFMC are not designated yet.

According to technical Annexes for the 2 lots, these annexes must be attached, duly completed, to the Tenderer's offer.

#### ARTICLE 5: Organization of delivery, installation, and commissioning reception

The delivery of equipment to the service premises, assembly, installation and commissioning are the responsibility of the Tenderers,

They must explicitly show the conditions of ground load, electrical power supply, average power, peak power absorbed, the tolerable temperature in each room (in particular the maximum tolerable variation gradient).

The Tenderer is deemed to know the places for having visited them. The visit may take place at his own discretion. To do this, it will be up to him to make a written request to the Hospital Centre or CDE, which will set up an appointment for him, as precise in section III.7.

If disturbances are observed when the equipment is put into service or immediately afterward, it is the holder's responsibility to take all useful measures to remedy them under conditions approved by the Hospital Center or, failing that, to provide proof recognized by an official body that these disorders are not attributable to the interventions made under its responsibility or to the equipment installed by it.

The mandatory tests and inspections for the commissioning of the equipment and the verification of their performance will be carried out before receipt according to the manufacturer's recommendations and the regulations in force.

The cost of these operations is the responsibility of the Tenderer.

CDE shall bear no responsibility over losses or damages of the procured products incurred during the performance period and before acceptance of said products. The supplier bears the responsibility to verify and certify that the goods they supply are in keeping with the conditions applicable to them.

The supplier commits to provide CDE with goods that will not be subject to manufacturing defects, that have not been exposed to contamination or to anything causing premature wear. Products supplied by the Supplier are covered by a warranty period.

It is therefore up to the supplier to ensure the products (mandatory).

The total quantity of the product should be well packed in compliance with best international and professional practices or rules.

Before each delivery, the supplier has to submit a copy of the delivery note and all the shipping documents detailed below BEFORE loading and shipping the goods, in order to get the formal agreement from CDE to deliver the goods.

For every consignment, the supplier shall always send a delivery note. Delivery slips shall necessarily bear the Contract Reference and / or Purchase order number, batch numbers, serial numbers if any, the full designation and quantities of the delivery.

Added to the delivery note, the selected supplier will also have to provide CDE with:

- a delivery notes and / or Packing List

- a commercial invoice

- a Certificate of Origin (if applicable)

- a Certificate of conformity or Certificates of Analysis (if applicable)

- other documents.

The Supplier commits to inform CDE of any constraint or specific regulation linked to the goods or service supply or to the country of importation

In case of international sea or air transport: After the agreement from the entity that placed the order, 3 original set + 1 copy of these documents will be sent by express courier within four

(4) working days after products have been loaded to the CDE entity that placed the order. A Bill of lading or Airway Bill (3 originals + 1 copy) fulfilled with all the required information has to be added to the documentation.

The supplier shall put in place, and communicate to CDE, their internal quality control system, if CDE deems it necessary for the guarantee of the supplier’s products.

The Supplier will inform CDE about all quality certifications, labels (NF, ISO, CE…) and internal quality process that may apply to its goods or services and will supply all official documents upon CDE request.

CDE reserves the right to verify or use the services of a third party of its choice to verify the implementation by the supplier of the quality control procedures laid down in the supplier’s quality control system.

###### Non-conformity of delivery

Should the quality or the condition of the products not satisfy CDE requirements at the moment of the preliminary inspection or delivery inspection, CDE reserves the right to demand:

- The delivery of products which conforms to the order. They will need to be replaced by the Tenderer at his/her own expenses. The replacement will be executed as soon as possible, at latest within fifteen (15) calendar days from the discovery of the non-compliance (from the issuing of a certificate of non-conformity). The replaced products will again be subject to the rules laid down in this contract, including the twelve (12) month guarantee (IF APPLICABLE)

- Or the immediate reimbursement of the payment

- Or the cancellation of the order and of the corresponding price.

#### ARTICLE 6: Upgrade

The contract being notified for several years, the homogeneous evolution of all the equipment is planned with this contract.

This evolution concerns both the hardware part, software and accessories.

For each of the products and software offered, the Tenderer will list the major hardware and/or software developments already implemented since the date of first commissioning.

It will imperatively be indicated the version of the software provided.

The Tenderer will specify which hardware and/or software developments are planned, detailing:

- the nature of the expected change,

- the methods of implementation,

- the planned publication date or the planned periodicity,

- the approximate additional implementation cost.

The Tenderer will also indicate which characteristics (technology, architecture, etc.) facilitate the subsequent development of the products offered.

The Tenderer will specify the terms of support for upgrades, hardware and software developments and undertakes to do so contractually.

The Tenderer will commit in particular to the tariffs and/or the minimum discount rates which will be practiced during the operation of the equipment.

In the event of technological development of its products, during the period of execution of the contract, the selected supplier may propose to substitute a new product for the old one. This substitution can only be effective after written acceptance by the manager of the establishment and without change in the unit price.

#### ARTICLE 7: Medical equipment fitting and Room layouts (loaded drawings)

The Tenderer has to provide the medical equipment that will be fitted according to room layout. In this tender, the specific room details will be provided later to recognize the equipment locations . And it will be enclosed and completed in the (ANNEX 7 - SFMC Presentation and Layouts)

### LIST OF ANNEXES:

ANNEX 1 - Pricing schedule Lot 1: Imaging equipment (To be signed, stamped and submitted)

ANNEX 2 - Pricing schedule Lot 2: Monitoring and Ventilation equipment (To be signed, stamped and submitted)

ANNEX 3 to 5: not applicable

ANNEX 6 - Biomedical tender schedule ANNEX 7 - SFMC Presentation and Layouts

ANNEX 8 - CDE Ethics Clauses (To be signed, stamped and submitted)

ANNEX 9 - Impartiality document (To be signed, stamped and submitted)

ANNEX 10 - Non-disclosure agreement (To be signed, stamped and submitted)

ANNEX 11- Checklist Tender documents submission (To be signed, stamped and submitted)

ANNEX 12 – Supply contract

# TECHNICAL SPECIFICATIONS

##### Technical Data Instructions

Technical data for each medical equipment specialty is divided into LOTS for specifications and requirements.

Note: For the submission of tenders, each candidate will propose a plan layout (if possible, in 3D) of the proposed equipment as well as the standard equipment of an operating room (table, respirator, mobile radio, etc.).

The layout will make it possible to check the ergonomics of the rooms by showing the effects of handling the domes and arms.

## Lot 1: Supply, installation, qualification and start-up of Imaging Equipment.

##### Object

In the present tender, the supplier will provide the Imaging equipment, as well as delivery, installation and commissioning, staff training, and maintenance services to the Sinjar site (SFMC - SINJAR FRENCH MEDICAL CENTRE).

The supplier will be  responsible for obtaining and paying for insurance coverage that protects the goods during transit until they are delivered to the named Sinjar -SFMC._ DAP Incoterm_

##### Duration

All equipment must be operational on the planned taking-over date of the SFMC works, planned at the beginning of 2024

The Supplier will be responsible to provide his detailed delivery, installation and training schedule which shall be validated by La Chaîne de l’Espoir.

##### Allotment

This tender is composed of a single lot for equipment and services. The following is a complete list of equipment from lot 1 with quantities:

| Equipment | Quantity |
| --- | --- |
| 1. Ultrasound Machine | 1 |
| 2. Echocardiography unit | 1 |
| 3. IMAGER, dry laser printer, floor standing | 1 |

##### Power requirement for the medical devices

Outlets and receptacles have the following specifications: 13A, 220V, 1-Ph, 2-P, 3-Wire. - 50/60 Hz. The medical equipment should therefore conform to these specifications.

##### This lot also includes the services below:

The lot include the procurement of a from supply to delivery to installation and commissioning. The services include:

- Preparatory work for the installation

- Supply, delivery, installation

- The restoration of all damage caused during the delivery and assembly of the devices.

- Assembly and commissioning

- Administrative procedures and on-site inspections and tests (regulatory or necessary for commissioning) including the consumables necessary for these tests, Guidelines, Quality Assurance, Software verification, periodic maintenance schedule

- 2 years warranty includes preventive and curative maintenance (travel, hand work and parts)

- Training of technicians and engineers. This training will include travel and accommodation if necessary. Following this training, engineers and technicians will be handling and commissioning the equipment and authorized to intervene on medical devices within the limits specified by this authorization.

- The documentation of use and certificate in English and Arabic

##### As part of the contract, each of the Tenderers must:

      - Define a phasing schedule for the commissioning of the various equipment proposed, with deadlines for delivery, installation and training.

      - Specify for each piece of equipment:

        - Environment conditions: evacuation, air-conditioning of the premises (temperature and hygrometry conditions, heat release of the equipment) and necessary extraction required per unit

        - Necessary power supply

        - Voltage (values, tolerance)

        - Frequency (values, tolerance)

        - Maximum power consumption

##### Object Standards and regulations

Standards and regulations

Medical devices, their components, and the services provided must comply with all applicable standards and regulations mentioned in this section on the date of contract signature, as well as those applicable at the time of final acceptance of equipment and services:

- ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU)

-	ISO/IEC 80601-2-13:20##.

- IEC 60601-1- General requirements (or equivalent).

-	ISO 10651-3.

##### CE marking

All equipment, peripherals, accessories and consumables will comply with the regulations relating to Directive 93/42/EEC.

The installation must be in accordance with standard and all the regulations in force (hygiene, machine safety, electrical safety, etc.).

Each medical device and each pressure device will be accompanied by CE declarations and CE type examination certificates with the limit of validity and recurrence of controls.

##### FDA clearance (Preferred) and ISO certificates

The required certificates to be provided.

##### Object Associated maintenance contract and guarantee:

Guarantee:

- The Tenderer’s offer will specify the exact duration of the guarantee (the minimum warranty period is 1 year).

- The starting point of the guarantee period is the date of commissioning and acceptance by the client of the equipment as stated in the decision to accept the equipment.

- Interventions lasting more than two working days carried out during the warranty period, postpone the term of the warranty accordingly

- The Tenderer will provide full maintenance coverage "parts, labor and travel" for the entire warranty period.

- The Tenderer will clearly specify, if necessary, the exclusions or the limits of guarantee.

##### Warranty maintenance:

Preventive and corrective maintenance and regulatory performance checks are carried out by the contractor during the specified 2 years warranty period, in accordance with the manufacturer's recommendations

The Tenderer will indicate the assumption of responsibility for the preventive maintenance during the duration of the guarantee.

##### Out-of-warranty maintenance:

The Tenderer will specify in his offer the elements relating to the support of out-of-warranty maintenance with an offer including:

- Training of the technical team in curative and preventive maintenance of equipment. The equipment and conditions of realization necessary for a support by the maintenance service of the hospital.

- labor, parts (excluding consumables) for support by the supplier's after-sales service, for maintenance from level 2.

The services are carried out in partnership with the technical staff of the Biomedical department of the hospital according to their availability. However, the hospital reserves the right to involve the contracting party if necessary. The Tenderers will indicate in his offer the amounts relating to this type of intervention.

##### Object Technical characteristics of lot 1 equipment

      - Technical characteristics expected for Ultrasound Machine (indicative quantity: 1):

| Equipment | TECHNICAL CHARACTERISTICS |
| --- | --- |
| Room code | EXAM, PREP ROOM: N1_OB_004 |
| Ultrasound Machine | REQUIREMENTS: CLINICAL APPLICATIONS : - Abdominal, - OB/GYN, - small parts, - urology , - vascular, - breast, - others .Type: On-platform ,CE /) PROBE TYPES, - MHz:Linear array( multi frequency ) :yes, - Convex array ( multi frequency ) : yes , - Endovaginal array ( multi frequency ) : yes , SCAN MODES: - B –mode( 2-D) :yes , - M –mode( 2-D) :yes ,3-D (automatic) :yes , - Harmonic tissue imaging :yes ), - DOPPLER:(Color Doppler imaging (CDI) :yes , - Power Doppler imaging (PDI) : yes , - Continuous wave, Pulsed wave:yes , - Duplex mode :yes ,Triplex mode :yes, All option should be permanent activation, - Image Display and Processing : - Automated image optimization :yes , - Gray scale level: ≥ 256 ,Image magnification :yes , - Monitor (in): ≥ 12 ,Maximum display ,mm :≥ 240 ,Preprocessing , Postprocesing , - Real-time image , - Frozen image , - SPLIT SCREEN , - ANALYSIS PACKAGES ,abdominal, neck ultrasound, breast, OB/GYN, small parts, urology and others ,FUNCTIONALITY (Digital calipers : yes), Techincal Spesifcation: - IMAGE STORAGE : HDD,DVD, - USB, - printer, - other should specified (Capacity, number of stored images: ≥ 250 GB HD ,Printer : Yes - ( B & W ) or better ,Cine :yes) - NUMBER OF USER-PROGRAMMABLE PROTOCOLS, FORMULAS , TABLES :yes , - Power supply : Medical Approved power supply board ,Line voltage: - 220/240VAC,50/60Hz, - Three pin G type plug ,Printer : yes , - Backup Battery :yes , - The equipment suitable for working in the climate condition in Iraq in terms of condition and humidity FDA:Yes ( FDA preferred optional : DICOM |

##### Technical characteristics expected for Echocardiography unit (indicative quantity: 1):

| Equipment | TECHNICAL CHARACTERISTICS |
| --- | --- |
| Room code | Recovery room: N1_SD_003 |
| Echocardiography unit | REQUIREMENTS: CLINICAL APPLICATIONS: - Cardiac (adult, pediatric, neonatal, fetal, TEE, intracardiac), - vascular, - stress echo, - others .Type: On-platform ,CE / FDA:Yes ( FDA preferred ), PROBE TYPES, - MHz: Phased array( multi frequency ) :yes, - Linear array ( multi frequency ) : yes , - SCAN MODES:(B –mode( 2-D) :yes , - M mode, yes. Scan modes - Harmonic tissue imaging :yes ), - DOPPLER:(Color Doppler imaging (CDI) :yes , - Tissue Doppler imaging (TDI) :yes, - Power Doppler imaging (PDI) : yes , - Continuous wave : yes, - Pulsed wave:yes ,Duplex mode :yes , - Triplex mode :yes, All option should be permanent activition , - Image Display and Processing : - Automated image optimization :yes , - Gray scale level: ≥ 256 ,Image magnification :yes , - Monitor (in): ≥ 12 ,Maximum display ,mm :≥ 240 ,Preprocessing , Postprocesing , - Real-time image , - Frozen image , - SPLIT SCREEN , - ANALYSIS PACKAGES, Cardiac, vascular ,user-programmable calculations, etc FUNCTIONALITY ( Digital calipers : yes) Techincal Spesfication:, IMAGE STORAGE : HDD,DVD, USB, printer, other should specified (Capacity, number of stored images: ≥ 250 GB HD ,Printer : Yes ( B & W ) or better ,Cine :yes) NUMBER OF USER-PROGRAMMABLE PROTOCOLS, FORMULAS , TABLES :yes ,Power supply : Medical Approved power supply board ,Line voltage: - 220/240VAC,50/60Hz, Three pin G type plug ,Printer : yes ,Backup Battery :yes ,The equipment suitable for working in the climate condition in Iraq in terms of condition and humidity Other Option: - ECG - DICOM |

      - Technical characteristics expected for IMAGER, dry laser printer, floor standing (indicative quantity: QTY 1):

| Equipment | TECHNICAL CHARACTERISTICS |
| --- | --- |
| Room code | FILM PROCESSING: N1_RA_012 |
| IMAGER, dry laser printer, floor standing | Rated frequency range: 50 to 60 Hz Input frequency range: 49 to 61 Hz Rated current: 1.4 - 0.7 A Voltage: 220V Cassette Size. 10 x 12 & 12 x 14 inch. Printing Resolution:. 320 dpi. Pixel Depth. 14 Bit. Power Consumption. 530 Watt. Voltage 240 V Capacity Of Supply Tray 100 Sheets 2 film trays UPS is an option |

## Lot 2: Supply, installation, qualification and start-up of Monitoring and Ventilation equipment.

##### Object

In the present tender, the supplier will provide the Operating room equipment, as well as delivery, installation and commissioning, staff training, and maintenance services to the Sinjar site (SFMC - SINJAR FRENCH MEDICAL CENTRE).

DAP Incoterm, the supplier is responsible for obtaining and paying for insurance

coverage that protects  the goods during transit until they are delivered to Sinjar -SFMC..

##### Duration

All equipment must be delivered and operational after the planned taking-over date of the SFMC works, planned at the beginning of 2024.

The Supplier will be responsible to provide his detailed delivery, installation and training schedule which shall be validated by La Chaîne de l’Espoir.

##### Allotment

This tender is composed of a single lot for equipment and services. The following is a complete list of equipment from lot 2 with quantities:

| Equipment | Quantity |
| --- | --- |
| - VENTILATOR, Adult - Pediatric | 1 |
| - MONITOR Fetal / Maternal,Antepartum | 1 |
| - MONITOR, Physiological, Central PACU/ICU | 1 |
| - MONITOR, Vital Signs, on mobile stand | 1 |
| - MONITOR, pulse oximeter | 2 |
| - Fetal Monitor Sonicaid | 1 |
| - Intensif care ventilator | 2 |
| - SYRINGE pump | 1 |
| - SUCTION MACHINE Electric mobile, OB/GYN | 1 |
| - INFUSION volumetric pump, single channel | 1 |
| - REGULATOR suction, adjustable, continous/intermittent, Suction bottles | 1 |
| - ELECTRONIC THERMOMETER, Rechargeable, ward | 1 |
| - Flowmeter, Air | 4 |
| - FLOWMETER, Oxygen 0-15 LPM | 1 |

##### Power requirement for the medical devices:

Outlets and receptacles have the following specifications: 13A, 220V, 1-Ph, 2-P, and 3-Wire. - 50/60 Hz. The medical equipment should therefore conform to these specifications

##### This lot also includes the services below:

The lot includes the procurement of a from supply to delivery to installation and commissioning. The services include:

- Preparatory work for the installation

- Supply, delivery, installation

- The restoration of all damage caused during the delivery and assembly of the devices.

- Assembly and commissioning

- Administrative procedures and on-site inspections and tests (regulatory or necessary for commissioning) including the consumables necessary for these tests.

- Warranty includes preventive and curative maintenance (travel, hand work and parts) for 2 years.

- Training of technicians and engineers. This training will include travel and accommodation if necessary. Following this training, engineers and technicians will be able to work on the equipment handling and commissioning the equipment and authorized to intervene on medical devices within the limits specified by this authorization.

- The documentation of use and certificate in English and Arabic

##### As part of the contract, each of the Tenderers must:

      - Define a phasing schedule for the commissioning of the various equipment proposed, with deadlines for delivery, installation and training.

      - Specify for each piece of equipment:

        - Environment conditions: evacuation, air-conditioning of the premises (temperature and hygrometry conditions, heat release of the equipment) and necessary extraction required per unit

        - Necessary power supply

        - Voltage (values, tolerance)

        - Frequency (values, tolerance)

        - Maximum power consumption

##### Object Standards and regulations

Standards and regulations

Medical devices, their components, and the services provided must comply with all applicable standards and regulations mentioned in this section on the date of contract signature, as well as those applicable at the time of final acceptance of equipment and services.

##### CE marking

All equipment, peripherals, accessories and consumables will comply with the regulations relating to Directive 93/42/EEC.

The installation must be in accordance with standard and all the regulations in force (hygiene, machine safety, electrical safety, etc.).

Each medical device and each pressure device will be accompanied by CE declarations and CE type examination certificates with the limit of validity and recurrence of controls.

##### FDA clearance (Preferred) and ISO certificates

The required certificates to be provided.

##### Object Associated maintenance contract and guarantee:

Guarantee:

- The Tenderer’s offer will specify the exact duration of the guarantee.

- The starting point of the guarantee period is the date of commissioning and acceptance by the client of the equipment as stated in the decision to accept the equipment.

- Interventions lasting more than two working days carried out during the 2 years warranty period, postpone the term of the warranty accordingly

- The Tenderer will provide full maintenance coverage "parts, labor and travel" for the entire warranty period.

- The Tenderer will clearly specify, if necessary, the exclusions or the limits of guarantee.

##### Warranty maintenance:

Preventive and corrective maintenance and regulatory performance checks are carried out by the contractor during the specified warranty period, in accordance with the manufacturer's recommendations

The Tenderer will indicate the assumption of responsibility for the preventive maintenance during the duration of the guarantee.

##### Out-of-warranty maintenance:

The Tenderer will specify in his offer the elements relating to the support of out-of-warranty maintenance with an offer including:

- Training of the technical team in curative and preventive maintenance of equipment. The equipment and conditions of realization are necessary for a support by the maintenance service of the hospital.

- labor, parts (excluding consumables) for support by the supplier's after-sales service, for maintenance from level 2.

The services are carried out in partnership with the technical staff of the Biomedical department of the hospital according to their availability. However, the hospital reserves the right to involve the contracting party if necessary. The Tenderers will indicate in his offer the amounts relating to this type of intervention.

##### Object Technical characteristics of lot 2 equipment

      - VENTILATOR, Adult - Pediatric (indicative quantity: QTY 1):

| Equipment | TECHNICAL CHARACTERISTICS |
| --- | --- |
| Room code | RESUSCITATION ROOM: N1_ED_008 |
| VENTILATOR, Adult - Pediatric | REQUIREMENTS - ventilator suitable for adult, pediatric, and newborn patients - mix modes (11 Ventilation modes) - self-test and sensor calibration. - Apnea / backup ventilation function required - Drive mechanism (Turbine)=0.5% - AFNOR (O2)+MA4) - Battery capacity - ISO ,CE certificate = Compulsary - FDA certificate |

      - MONITOR Fetal / Maternal,Antepartum (indicative quantity: QTY 1):

| Equipment | TECHNICAL CHARACTERISTICS |
| --- | --- |
| Room code | Operating room 2 obstetric: N1_SD_014 |
| MONITOR Fetal / Maternal,Antepartum | REQUIREMENTS - 10.2" folding color widescreen - Multi-crystal pulsed wave transducer - FM marker, automatic fetal movement detection - Signal overlap verification - 12 hours data storage and review - Specifications - Physical Weight & Dimensions - Dimensions: 347 × 330 × 126 mm - Weight: Approx. 6 kg - Performance - Display: 10.2" color TFT-LCD - Resolution: 800 × 480 - Paper: Z-fold, thermosensitive - Recorder: US/International Standard; 1/2/3 cm/min real-time printing speed; fast print speed (stored traces) up to 25 mm/sec- ISO ,CE certificate = Compulsary - FDA certificate |

      - MONITOR, Physiological, Central PACU/ICU (indicative quantity: QTY 1):

| Equipment | TECHNICAL CHARACTERISTICS |
| --- | --- |
| MONITOR, Physiological, Central PACU/ICU | REQUIREMENTS Suitable for Adult, Pediatric, and Neonate Patient Support. With the following monitored parameters: ECG, HR (Heart rate), SpO2 (Pulse Oximetry), PR (Pulse Rate), RESP (Respiration), NIBP (Non-Invasive Blood Pressure), TEMP (Temperature), IBP. ECG 5 leads. Heart rate (HR): measuring range up to 250 bpm. Pulse Oximetry (SpO2): 1 to 100%, pulse rate (PR): up to 240 bpm. Respiration (RESP): measuring range 0 to 120 bpm, with apnea detection adjustable by the user. Non-invasive blood pressure (NIBP): manual, continuous (up to 5 min) or STAT mode. With an overpressure protection for the NIBP cuff. Temperature (TEMP): 1 channel, range up to 45°C, accuracy ± 0.1° C. With adjustable, visual and audible alarms, with 3 levels of priority (high, medium, low). LCD/TFT display, with simultaneous display of minimum 4 traces and 8 vital parameters. With a Multilanguage interface. English language must be included as a minimum. Data storage 24 hours minimum. Export via SD card or USB port. Connectable to network, via Ethernet or wireless / ISO ,CE certificate = Compulsary FDA certificate |

      - MONITOR, Vital Signs, on mobile stand (indicative quantity: QTY 1):

| Equipment | TECHNICAL CHARACTERISTICS |
| --- | --- |
| MONITOR, Vital Signs, on mobile stand | REQUIREMENTS Adult SpO2 Sensor Standard plus Bonus Large NIBP Cuffs 3 Lead ECG Surface Temp Li-ion Battery (up to 10 hrs) Includes Stand Economic Patient Monitor 8” colour TFT LCD Screen (touch screen option) Portable, ergonomic, light weight, sturdy design Suitable for adult, pediatric & neonatal patients Flexible parameters configuration for different clinical environments Standard Configuration 3 Lead ECG, NIBP, Spo2, Temp Spot-check and continuous monitoring mode with alarms Graphical & tabular trend review with customisable size & colour settings Rechargeable Li-ion Battery (up to 10 hours) 48 hours holographic wave review for each patient/ ISO ,CE certificate = Compulsary FDA certificate |

      - MONITOR, pulse oximeter (indicative quantity: QTY 2):

| Equipment | TECHNICAL CHARACTERISTICS |
| --- | --- |
| MONITOR, pulse oximeter | Requirements 5 Standard parameters: ECG, RESP, NIBP, SPO2, 2-TEMP, PR/HR. 12.1"high resolution color TFT LCD display. Real time S-T segment analysis, pace-maker detection. Patient info input management function Multi-lead ECG waveforms display in phase. Large volume of tabular and graphic trends information storage and easy to recall. Capture dynamic waveforms. Efficient resistance to interference of defibrillator and HF knife. UP to 4 hours working capacity of built-in rechargeable battery. Networking capacity/ ISO ,CE certificate = Compulsary FDA certificate |

      - Fetal Monitor Sonicaid (indicative quantity: QTY1):

| Equipment | TECHNICAL CHARACTERISTICS |
| --- | --- |
| Fetal Monitor Sonicaid | Requirements FDA CLEARANCE and /or CE MARK (MDD) : yes preferred (FDA) TYPE/CONFIGURATION : tabletop CONSTRUCTION :• ABS, Polycarbonate, stainless steel, or other. FHR DISPLAY RANGE, bpm :• ≤ 240 PROBE Frequency, MHz : ≤ 10 Output, mW/cm² : ≤ 22 Diameter, mm : ≤40 Length, mm : • ≥100 FREQUENCY OUTPUT Audio jack, Hz : ≤ 3000 Speaker, Hz : ≤ 3000 Internal Speaker : yes Display : LCD, LED or better. Display variables : FHR, other. CONTROLS :• Volume, power on/off, (Other should be Specified). POWER REQUIREMENTS :• Power input to be 220-240VAC, 50 – 60 Hz. Line, VAC : Three Pin G Type Plug Battery : yes Operating time, hr : ≥ 5 Recharge time, hr. : ≤ 4 (if Rechargeable battery used) Charger) 220-240VAC, 50 – 60 Hz ( :• Yes (for Rechargeable battery type Low-battery indicator : yes STANDARD ACCESSORIES/ ISO ,CE certificate = Compulsary FDA certificate • Additional Probe • Headphone • Ac adapter (if need for Charger) • soft carry bag. • Gel (60g) (Qty.) ≥ 2 ,,• The supplier should be confirmed that: The Equipment is suitable for work in climate conditions of Iraq in terms of temperature and humidity. |

      - Intensif care ventilator (indicative quantity: QTY 2):

| Equipment | TECHNICAL CHARACTERISTICS |
| --- | --- |
| Intensive care ventilator | Requirements -ventilator suitable for adult, pediatric, and newborn patients mix modes (11 modes) Ventilator self-test and sensor calibration. Apnea / backup ventilation function required Drive mechanism (Turbine) BS (O2)+MA4) - Battery capacity ISO ,CE certificate FDA certificate |

      - SYRINGE pump (indicative quantity: QTY 1):

| Equipment | TECHNICAL CHARACTERISTICS |
| --- | --- |
| SYRINGE pump | Requirements - 20-, 50- and 60-ml syringes. - Automatic detection of syringe size and proper fixing. - Anti-bolus system - Self-check carried out on powering on ISO ,CE certificate FDA certificate |

      - SUCTION MACHINE Electric mobile, OB/GYN (indicative quantity: QTY 1):

| Equipment | TECHNICAL CHARACTERISTICS |
| --- | --- |
| Trolley-mounted oxygen therapy system | Requirements CONFIGURATION Mobile on cart with castors PUMP TYPE : Piston, or Diaphragm, Others to be specified Number of pumps : ≥ 1 VACUUM LIMIT, mm Hg, (cmH2O) : ≥ 500 , (679.75) Free airflow, L/min : ≥25 VACUUM CONTROL : YES COLLECTION CANISTER(S) : Reusable &Disposable Number : ≥2 Capacity limit line, mL : ≥2000 Overflow protection YES FILTER TYPE Bacterial, hydrophobic STANDARD ACCESSORIES (Included in Warranty) : • Collection bottle. • Filters. • Cart. Electrical Ratings : - Power Supply -220/240VAC - Rated Frequency 50/60Hz Environmental The supplier should be confirmed that: The Equipment is suitable for work in climate conditions of Iraq in terms of temperature and humidity. Note : glass – cylinder / ISO ,CE certificate = Compulsary FDA certificate |

      - INFUSION volumetric pump, single channel (indicative quantity: QTY 1):

| Equipment | TECHNICAL CHARACTERISTICS |
| --- | --- |
| INFUSION volumetric pump, single channel | Requirements -programmable range at least from 0.1 to 200 ml/hr -Bolus rate approx. 500 ml -occlusion pressure from 300, 500 and 900 mmHg. -Comprehensive alarm package -It should be open system. ISO ,CE certificate FDA certificate |

      - REGULATOR suction, adjustable, continous/intermittent, Suction bottles (indicative quantity: QTY1):

| Equipment | TECHNICAL CHARACTERISTICS |
| --- | --- |
| REGULATOR suction, adjustable, continous/intermittent, Suction bottles | Requirements Interval of manometer, level of bottles, if antibacterian filter, pressure and probe FDA CLEARANCE and /or CE MARK (MDD) : yes preferred (FDA) TYPE/CONFIGURATION : tabletop CONSTRUCTION :• ABS, Polycarbonate, stainless steel, or other. FHR DISPLAY RANGE, bpm :• ≤ 240 PROBE Frequency, MHz : ≤ 10 Output, mW/cm² : ≤ 22 Diameter, mm : ≤40 |

      - ELECTRONIC THERMOMETER, Rechargeable, ward (indicative quantity: QTY 1):

| Equipment | TECHNICAL CHARACTERISTICS |
| --- | --- |
| ELECTRONIC THERMOMETER, Rechargeable, ward | Requirements Temperature measurement range 32 – 43 °C (minimum guaranteed). Accuracy ± 0.1°C in the range 35 – 41 °C. Graduation 0.3°C or better. Ready-to-use after switch-on within 10 s. with battery /ISO ,CE certificate = Compulsary FDA certificate |

      - Flowmeter, Air (indicative quantity: QTY 4):

| Equipment | TECHNICAL CHARACTERISTICS |
| --- | --- |
| Flowmeter, Air | Requirements Flow Range 0 to 300 SLPM of Air Flow Accuracy 2% of reading or 0.05 SLPM, whichever is greater, at 70 °F and 14.7 psia Flow Sample Rate 50 milliseconds Temperature Range 0 to 50 °C (0 to 122 °F) Accuracy ±1 °C (±1.8 °F) at flows >1 SLPM Atmospheric Pressure Range 15 to 115 kPa (2.17 to 16.68 psia) Atmospheric Pressure Accuracy ±1.5% F.S./ ISO ,CE certificate = Compulsary FDA certificate |

      - FLOWMETER, Oxygen 0-15 LPM (indicative quantity: QTY 1):

| Equipment | TECHNICAL CHARACTERISTICS |
| --- | --- |
| FLOWMETER, Oxygen 0-15 LPM | Requirements 0-15 LPM, accuracy 10%, tube gradation max 0.5 LPM at low flow, max 1 LPM high flow Outlet connector for 6 mm barbed connection to medical air tubing included. Certificate of calibration and inspection/ ISO ,CE certificate = Compulsary FDA certificate |